## Table 2
Table 2. Analytic cross-reactivity of “in-house” real-time PCR assay for nine STIs
STT | Micro-organism | Name/origin of strain or its DNA control | Master mix 1 | Master mix 2 | Master mix 3
|  |  | (GV/NG/CT) | (HSV-1/CA/TV) | (MH/HSV-2/MG)
Targeted 9 STIs
1 | G. vaginalis | ATCC 14018 | 20.22 | Negative | Negative
|  | PS GV DNA Control | 16.43 | Negative | Negative
2 | N. gonorrhoeae | ATCC 19424 | 27.93 | Negative | Negative
|  | BM NG DNA Control | 28.53 | Negative | Negative
3 | C. trachomatis | Amplirun Chlamydia trachomatis DNA Control | 30.72 | Negative | Negative
|  | BM CT DNA Control | 26.74 | Negative | Negative
4 | HSV-1 | Amplirun Herpes Simplex 1 DNA Control | Negative | 23.75 | Negative
|  | BM HSV-1 DNA Control | Negative | 15.9 | Negative
5 | C. albicans | ATCC14053 | Negative | 29.78 | Negative
|  | BM CA DNA Control | Negative | 30.39 | Negative
6 | T. vaginalis | Amplirun Trichomonas vaginalis DNA Control | Negative | 24.12 | Negative
|  | PS TV DNA Control | Negative | 29.3 | Negative
7 | M. hominis | Amplirun Mycoplasma hominis DNA Control | Negative | Negative | 15.89
|  | PS MH DNA Control | Negative | Negative | 33.83
8 | HSV-2 | Amplirun Herpes Simplex 2 DNA Control | Negative | Negative | 20.39
|  | BM HSV-2 DNA Control | Negative | Negative | 24.13
9 | M. genitalium | Amplirun Mycoplasma Genitalium DNA Control | Negative | Negative | 24.34
|  | BM MG DNA Control | Negative | Negative | 21.02
Non-targeted STI: HPV genotypes
1 | HPV genotype 16 | BM HPV16 DNA Control | Negative | Negative | Negative
2 | HPV genotype 18 | BM HPV18 DNA Control | Negative | Negative | Negative
3 | HPV genotype 31 | BM HPV31 DNA Control | Negative | Negative | Negative
4 | HPV genotype 33 | BM HPV33 DNA Control | Negative | Negative | Negative
5 | HPV genotype 35 | BM HPV35 DNA Control | Negative | Negative | Negative
6 | HPV genotype 39 | BM HPV39 DNA Control | Negative | Negative | Negative
7 | HPV genotype 45 | BM HPV45 DNA Control | Negative | Negative | Negative
8 | HPV genotype 51 | BM HPV51 DNA Control | Negative | Negative | Negative
9 | HPV genotype 52 | BM HPV52 DNA Control | Negative | Negative | Negative
10 | HPV genotype 56 | BM HPV56 DNA Control | Negative | Negative | Negative
11 | HPV genotype 58 | BM HPV58 DNA Control | Negative | Negative | Negative
12 | HPV genotype 59 | BM HPV59 DNA Control | Negative | Negative | Negative
13 | HPV genotype 66 | BM HPV66 DNA Control | Negative | Negative | Negative
14 | HPV genotype 68 | BM HPV68 DNA Control | Negative | Negative | Negative
Non-STIs bacteria
1 | B. subtilis | VTCC 11010 | Negative | Negative | Negative
2 | B. clausii | VTCC 11111 | Negative | Negative | Negative
3 | L. acidophilus | VTCC 12284 | Negative | Negative | Negative
4 | L. casei | VTCC 10411 | Negative | Negative | Negative
5 | L. plantarum | VTCC 10936 | Negative | Negative | Negative
6 | L. rhamnosus | VTCC 11354 | Negative | Negative | Negative
7 | E. coli | VTCC 12272 | Negative | Negative | Negative

## Table 3
Table 3. Performance of “in-house” multiplex real-time PCR assay for STI identification comPred to real-time PCR commercial kits used in hospitals
STT | Code | Full name | Born | Allplex Vaginitis Screening Assay |  |  | Allplex CT/NG/MG/TV Assay |  |  |  | Comercial kit Result | "In house" kit result | In house kit | [5 empty] | Comercial kit |  |  | In house kit
|  |  |  | CA (Ct) | GV (Qt) | TV (Ct) | NG (Ct) | CT (Ct) | TV (Ct) | MG (Ct) |  |  | GV (Ct) | NG (Ct) | CT (Ct) | CA (Ct) | TV (Ct) | MG (Ct) | Abr Name Pthogens | n | % | Abr Name Pthogens | n | %
1 | HOG _04 | Pham T.V.A | 1981 |  | 2.17 | [5 empty] | GV | GV | 31.9 | [5 empty] | GV | 25 | 20.33% | GV | 25 | 20.33%
2 | HOG _05 | Nguyen T.K.C | 1997 | 27.33 | 6.82 | [5 empty] | GV+CA | GV+CA | 16.92 |  |  | 27.15 |  |  | TV | 15 | 12.2% | TV | 15 | 12.2%
3 | HOG _06 | Nguyen M.H | 1999 |  | 3.7 | [5 empty] | GV | GV | 25.45 | [5 empty] | NG | 3 | 2.439% | NG | 3 | 2.439%
4 | HOG _07 | Pham T.V | 1989 | [7 empty] | Negative | Negative | [6 empty] | MG | 8 | 6.504% | MG | 8 | 6.504%
5 | HOG _08 | Tran T.H.N | 2001 | 28.4 | 7.02 | [5 empty] | GV+CA | GV+CA | 15.37 |  |  | 28.52 |  |  | GV+CA | 23 | 18.7% | GV+CA | 24 | 19.51%
6 | HOG _09 | Pham T.H | 1990 | 25.44 | 7.62 | [5 empty] | GV+CA | GV+CA | 14.56 |  |  | 2022-02-28 |  |  | GV+CT | 15 | 12.2% | GV+CT | 16 | 13.01%
7 | HOG _10 | Tran T.H | 1994 |  | 0.33 | [5 empty] | GV | GV | 35.52 | [5 empty] | CT+NG | 5 | 4.065% | CT+NG | 5 | 4.065%
8 | HOG _12 | Nguyen T.H | 1999 | 27.43 | 7.44 |  |  | 18.6 |  |  | GV+CT+CA | GV+CT+CA | 15.51 |  | 20.28 | 26.56 |  |  | GV+TV | 4 | 3.252% | GV+TV | 4 | 3.252%
9 | HOG _13 | Quach T.N.A | 2002 | 23.84 | 3.37 | [5 empty] | GV+CA | GV+CA | 27.42 |  |  | 27.77 |  |  | GV+MG | 1 | 0.813% | GV+MG | 1 | 0.813%
10 | HOG _16 | Hoang T.T.H | 1992 | 28.75 | 6.66 | [5 empty] | GV+CA | GV+CA | 17.22 |  |  | 30.20 |  |  | GV+NG | 1 | 0.813% | GV+NG | 1 | 0.813%
11 | HOG _17 | Le T.H | 1990 |  | 7.2 | [5 empty] | GV | GV | 16,45 | [5 empty] | GV+CT+CA | 9 | 7.317% | GV+CT+CA | 8 | 6.504%
12 | HOG _18 | Ngo T.H.H | 1988 | 24.92 | 7.53 | [5 empty] | GV+CA | GV+CA | 15.28 |  |  | 27.67 |  |  | GV+CT+MG | 3 | 2.439% | GV+CT+MG | 2 | 1.626%
13 | HOG _19 | Nguyen T.N.Q | 2002 | 27.13 | 7.31 | [5 empty] | GV+CA | GV+CA | 14.69 |  |  | 29.46 |  |  | GV+CA+MG | 1 | 0.813% | GV+CA+MG | 1 | 0.813%
14 | HOG _20 | Tran T.H | 1989 |  | 6.36 | [5 empty] | GV | GV | 17,75 | [5 empty] | GV+CT+TV | 2 | 1.626% | GV+CT+TV | 2 | 1.626%
15 | HOG _21 | Nguyen T.Q | 1997 | 28.83 | 6.96 | [5 empty] | GV+CA | GV+CA | 2022-03-16 |  |  | 27.46 |  |  | GV+NG+CT+CA | 3 | 2.439% | GV+NG+CT+CA | 3 | 2.439%
16 | HOG _22 | Tran T.V | 1965 |  | 6.79 | [5 empty] | GV | GV | 17,46 | [5 empty] | Negative | 5 | 4.065% | Negative | 5 | 4.065%
17 | HOG _23 | Ninh T.V.A | 1987 | 28.86 | 6.77 | [5 empty] | GV+CA | GV+CA | 17.72 |  |  | 29.95 |  |  | Total | 123 | 100% | Total | 123 | 100%
18 | HOG _24 | Trieu T.C | 1996 | 25.86 | 4.55 | [5 empty] | GV+CA | GV+CA | 24.71 |  |  | 28.88
19 | HOG _25 | Le T.H.V | 1981 |  | 0.92 | [5 empty] | GV | GV | 36.14 | [5 empty] | Abr Name Pthogens | Positive | Negative | Abr Name Pthogens | Positive | Negative
20 | HOG _36 | Nguyen T.T.T | 1977 |  | 1.41 | [5 empty] | GV | GV | 36.2 | [5 empty] | GV | 87 | 5 | GV | 87 | 5
21 | HOG _38 | Tran T.O | 1990 |  | 4.48 |  |  | 21.17 |  |  | GV+CT | GV+CT | 2022-02-22 |  | 22.73 |  |  |  | NG | 12 | 88 | NG | 12 | 88
22 | HOG _40 | Nguyen T.H | 1989 |  | 0.7 | [5 empty] | GV | GV | 35.83 | [5 empty] | CT | 37 | 60 | CT | 36 | 61
23 | HOG _49 | Lan N.Q | 2000 |  | 4.75 | [5 empty] | GV | GV | 23.96 | [5 empty] | HSV-1 | 16 | 76 | HSV-1 | 16 | 76
24 | HOG _60 | Vu H.H | 1972 |  | 2.72 | 17.97 |  |  | 20.76 |  | GV+TV | GV+TV | 29.81 |  |  |  | 24.48 |  | HSV-2 | 12 | 80 | HSV-2 | 12 | 80
25 | HOG _66 | Nong T.N.A | 1992 | [7 empty] | Negative | Negative | [6 empty] | CA | 36 | 56 | CA | 36 | 56
26 | HOG _68 | Bui T.T | 2000 | 24.91 | 6.85 |  | 25.87 | 23.99 |  |  | GV+NG+CT+CA | GV+NG+CT+CA | 15.44 | 2022-02-25 | 2022-10-25 | 26.77 |  |  | TV | 21 | 86 | TV | 21 | 86
27 | HOG _69 | Trinh Q.D | 1999 |  | 6.93 |  |  | 20.86 |  | 42.43 | GV+CT+MG | GV+CT | 15.77 |  | 21.22 |  |  |  | MH | 15 | 90 | MH | 15 | 90
28 | HOG _74 | Nguyen T.P.T | 1992 | 24 | 7.55 |  | 35.17 | 25.19 |  |  | GV+NG+CT+CA | GV+NG+CT+CA | 14.35 | 33.19 | 26.81 | 2022-06-26 |  |  | MG | 5 | 87 | MG | 4 | 88
29 | HOG _85 | Dinh T.T | 1977 | [7 empty] | Negative | Negative | [6 empty] | Total | 241 | 628 | Total | 239 | 630
30 | HOG _91 | Le T.T.H | 1991 | [7 empty] | Negative | Negative
31 | HOG _96 | Bui A.H | 1978 | 24.09 | 7.12 | [5 empty] | GV+CA | GV+CA | 15.11 |  |  | 27.06
32 | HOG _98 | Ninh T.V | 1978 | 24.52 | 7.54 | [5 empty] | GV+CA | GV+CA | 14.27 |  |  | 27.88
33 | HOG _99 | Do T.T.L | 1986 | 28.83 | 6.89 | [5 empty] | GV+CA | GV+CA | 15.71 |  |  | 29.29
34 | HOG _101 | Luong T.H | 1999 | 24.11 | 3.6 | [5 empty] | GV+CA | GV+CA | 26.41 |  |  | 26.92
35 | HOG _102 | Le T.L | 1994 |  | 6.72 | [5 empty] | GV | GV | 15.49
36 | HOG _106 | Nguyen T.P.C | 1994 | 24.91 | 7.51 |  |  | 30.42 |  |  | GV+CT+CA | GV+CT+CA | 14.54 |  | 29.61 | 28.1
37 | HOG _107 | Mai T.H.L | 1997 | 25.97 | 6.75 | [5 empty] | GV+CA | GV+CA | 16.26 |  |  | 26.37
38 | HOG _108 | Can T.H | 1996 | 32.5 | 7.64 |  |  | 19.38 |  |  | GV+CT+CA | GV+CT+CA | 14.1 |  | 19.96 | 33.31
39 | HOG _126 | Do B.N | 2003 |  | 3.13 |  |  | 20.85 |  |  | GV+CT | GV+CT | 27.82 |  | 24.4
40 | HOG _152 | Nguyen T.T | 1995 |  | 6.9 | [5 empty] | GV | GV | 15.89
41 | HOG _154 | Nguyen T.T.T | 1985 |  | 0.91 | [5 empty] | GV | GV | 35.06
42 | HOG _186 | Do T.T.T | 1976 | [7 empty] | Negative | Negative
43 | HOG _193 | Do T.T | 1993 | 34.38 | 6.2 | [5 empty] | GV+CA | GV+CA | 17.47 |  |  | 34.17
44 | HOG _197 | Vu T.H | 1978 |  | 5.4 | [5 empty] | GV | GV | 20.09
45 | HOG _201 | Le T.H | 1992 |  | 1.09 |  |  | 20.68 |  |  | GV+CT | GV+CT | 19.58 |  | 23.36
46 | HOG _221 | Dang P.T | 2001 |  | 6.49 |  |  | 38.17 |  |  | GV+CT | GV+CT | 17.32 |  | 37.32
47 | HOG _225 | Do T.T | 1992 | 25.46 | 4.1 | [5 empty] | GV+CA | GV+CA | 24.58 |  |  | 27.91
48 | HOG _230 | Bui T.L.Q | 1996 |  | 5.28 | 15.52 |  | 20.67 | 17 |  | GV+CT+TV | GV+CT+TV | 31.01 |  | 21.52 |  | 23.57
49 | HOG _238 | Nguyen V.A | 1988 |  | 5.61 | [5 empty] | GV | GV | 20.73
50 | HOG _264 | Tan T.D | 2004 |  | 2.29 |  |  | 20.48 |  | 37.75 | GV+CT+MG | GV+CT+MG | 34.38 |  | 21.17 |  |  | 36.33
51 | HOG _265 | Hoang T.T | 1995 |  | 6.58 | [5 empty] | GV | GV | 15.81
52 | HOG _266 | Nguyen T.H | 1997 |  | 2.62 |  |  | 18.8 |  |  | GV+CT | GV+CT | 27.01 |  | 17.35
53 | HOG _267 | Pham N.Q | 1991 | 23.77 | 7.33 | [5 empty] | GV+CA | GV+CA | 15.38 |  |  | 27.38
54 | HOG _284 | Nguyen T.S | 1968 |  | 5.76 |  |  |  | 19.61 |  | GV+TV | GV+TV | 19.94 |  |  |  | 29.3
55 | HOG _293 | Le T.H.A | 2001 |  | 5.3 | [5 empty] | GV | GV | 21.59
56 | HOG _295 | Lo T.H | 2002 |  | 5.03 |  |  | 22.33 |  |  | GV+CT | GV+CT | 23.3 |  | 28.86
57 | HOG _313 | Nguyen T.T | 1987 |  | 6.88 | [5 empty] | GV | GV | 16.16
58 | HOG _322 | Vu T.M | 2001 | 26.25 | 6.18 | [5 empty] | GV+CA | GV+CA | 19.21 |  |  | 29.72
59 | HOG _326 | Dang M.H | 1988 |  | 6.72 | [5 empty] | GV | GV | 16.1
60 | HOG _333 | Bui T.T | 2000 | 23.93 | 5.93 |  | 23.79 | 27.85 |  |  | GV+NG+CT+CA | GV+NG+CT+CA | 19.41 | 24.45 | 31.47 | 29.9
61 | HOG _334 | Ngo T.T.N | 1999 | 24.97 | 1.57 | [5 empty] | GV+CA | GV+CA | 32.82 |  |  | 29.78/30.39
62 | HOG _345 | Phung T.T | 1983 |  | 2.76 |  |  | 20.88 |  |  | GV+CT | GV+CT | 28.93 |  | 25.9
63 | HOG _349 | Le T.T.H | 1986 | 30.9 | 6.78 | [5 empty] | GV+CA | GV+CA | 15.59 |  |  | 29.44
64 | HOG _350 | Nguyen T.C.P | 1999 | 31.46 | 4.86 |  |  | 23.8 |  |  | GV+CT+CA | GV+CT+CA | 22.31 |  | 28.29 | 36.11
65 | HOG _351 | Phung L.A | 1978 |  | 6.58 | [5 empty] | GV | GV | 15.62
66 | HOG _353 | Vu T.T | 1995 |  | 4.77 |  |  |  |  | 30.95 | GV+MG | GV+MG | 16.77 |  |  |  |  | 33.41
67 | HOG _354 | Tran T.N | 1990 |  | 3.03 | 16.71 |  |  | 18.66 |  | GV+TV | GV+TV | 23.06 |  |  |  | 25.59
68 | HOG _356 | Duong T.H.L | 2001 | 23.45 | 6.54 |  |  |  |  | 28.6 | GV+CA+MG | GV+CA+MG | 16.82 |  |  | 26 |  | 32.11
69 | HOG _361 | Lo H.T | 1995 |  | 4.4 |  |  | 22.88 |  | 38.33 | GV+CT+MG | GV+CT+MG | 24.16 |  | 24.09 |  |  | 35.68
70 | HOG _365 | Duong T.L | 1977 |  | 0.93 | [5 empty] | GV | GV | 35.8 |  |  | -
71 | HOG _366 | Vu T.M.T | 2001 | 24.3 | 6.81 |  |  | 30.96 |  |  | GV+CT+CA | GV+CT+CA | 17 |  | 32.03 | 26.95
72 | HOG _390 | Tran T.H | 1991 |  | 6.51 |  |  | 20.32 |  |  | GV+CT | GV+CT | 17.28 |  | 20.76
73 | HOG _394 | Cu T.K.H | 1997 | 28.65 | 6.35 |  |  | 40.75 |  |  | GV+CT+CA | GV+CA | 19.15 |  |  | 28.71
74 | HOG _403 | Dao T.D | 1986 |  | 6.79 |  |  | 25.8 |  |  | GV+CT | GV+CT | 15.39 |  | 26.79
75 | HOG _414 | Vu B.P | 2004 |  | 5.77 | 16.44 |  | 24.19 | 18.16 |  | GV+CT+TV | GV+CT+TV | 18.89 |  | 25.01 |  | 24.83
76 | HOG _416 | Tran T.N.H | 2003 | 23.69 | 7.41 |  |  | 25.03 |  |  | GV+CT+CA | GV+CT+CA | 14.42 |  | 25.72 | 27.01
77 | HOG _419 | Vu T.Q | 1990 |  | 1.17 | [5 empty] | GV | GV | 35.65
78 | HOG _422 | Nguyen T.H | 1972 |  | 5.51 | 15.46 |  |  | 17.34 |  | GV+TV | GV+TV | 19.38 |  |  |  | 22.81
79 | HOG _433 | Le T.H | 1986 |  | 1.44 | [5 empty] | GV | GV | 32.49
80 | HOG _434 | Le T.N | 2002 |  | 7.11 | [5 empty] | GV | GV | 14.43
81 | HOG _435 | Do T.H.L | 2002 | 27.57 | 6.05 |  |  | 21.38 |  |  | GV+CT+CA | GV+CT+CA | 18.41 |  | 21.63 | 29.06
82 | HOG _440 | Tran T.B.T | 1987 |  | 7.06 | [5 empty] | GV | GV | 14.8
83 | HOG _441 | Vu T.D.Q | 1999 | 26.16 | 7.12 | [5 empty] | GV+CA | GV+CA | 15.67 |  |  | 29.01
84 | HOG _442 | Thai T.L.C | 2001 |  | 6.81 |  |  | 23.23 |  |  | GV+CT | GV+CT | 14.91 |  | 23.63
85 | HOG _447 | Trinh K.H | 2001 |  | 4.11 |  |  | 24.79 |  |  | GV+CT | GV+CT | 24.47 |  | 24.84
86 | HOG _466 | Nguyen T.P.L | 2003 |  | 1.61 |  | 25.46 |  |  |  | GV+NG | GV+NG | 30.99 | 23.96
87 | HOG _468 | Khuat N.H | 1999 |  | 1.43 |  |  | 23.85 |  |  | GV+CT | GV+CT | 33.41 |  | 27.45
88 | HOG _481 | Doan T.N.A | 2003 | 26.35 | 2.2 |  |  | 24.47 |  |  | GV+CT+CA | GV+CT+CA | 29.96 |  | 25.28 | 29.23
89 | HOG _484 | Kieu T.H | 1989 | 28.68 | 6.78 | [5 empty] | GV+CA | GV+CA | 15.04 |  |  | 29.55
90 | HOG _487 | Nguyen T.L.D | 1991 |  | 6.82 |  |  | 23.32 |  |  | GV+CT | GV+CT | 14.7 |  | 23.39
91 | HOG _493 | Nguyen H.T | 1990 |  | 3.4 |  |  | 25.36 |  |  | GV+CT | GV+CT | 27.78 |  | 29.85
92 | HOG _528 | Hoang T.H | 1993 |  | 2.28 |  |  | 23.63 |  |  | GV+CT | GV+CT | 31.39 |  | 25.33
93 | BM-TV_09 | Nguyen N. K. N. | 1994 |  |  | 16.11 |  |  |  |  | TV | TV |  |  |  |  | 22.23
94 | BM-TV_10 | Nguyen T. L. G. | 1988 |  |  | 20.78 |  |  |  |  | TV | TV |  |  |  |  | 24.39
95 | BM-TV_11 | Vu D. T. L. | 2000 |  |  | 15.18 |  |  |  |  | TV | TV |  |  |  |  | 21.16
96 | BM-TV_13 | Hoang H. D. | 1977 |  |  | 20.24 |  |  |  |  | TV | TV |  |  |  |  | 26.24
97 | BM-TV_14 | Phan N. N. T. | 1997 |  |  | 26.07 |  |  |  |  | TV | TV |  |  |  |  | 27.29
98 | BM-TV_15 | Dao T. K. | 1994 |  |  | 17.95 |  |  |  |  | TV | TV |  |  |  |  | 23.04
99 | BM-TV_16 | Dang G. H. | 1998 |  |  | 19.8 |  |  |  |  | TV | TV |  |  |  |  | 24.91
100 | BM-TV_17 | Tran N. M. A. | 2001 |  |  | 18.03 |  |  |  |  | TV | TV |  |  |  |  | 22.44
101 | BM-TV_18 | Hoang G. H. | 1984 |  |  | 19.45 |  |  |  |  | TV | TV |  |  |  |  | 24.88
102 | BM-TV_19 | Nguyen V. D. P. | 1995 |  |  | 16.39 |  |  |  |  | TV | TV |  |  |  |  | 21.24
103 | BM-TV_20 | Bui N. L. A. | 1991 |  |  | 19.87 |  |  |  |  | TV | TV |  |  |  |  | 24.28
104 | BM-TV_21 | Truong B. T. | 1992 |  |  | 18.5 |  |  |  |  | TV | TV |  |  |  |  | 25.43
105 | BM-TV_22 | Pham M. P. | 1985 |  |  | 19.57 |  |  |  |  | TV | TV |  |  |  |  | 25.41
106 | BM-TV_23 | Hoang B. A. | 1998 |  |  | 20.15 |  |  |  |  | TV | TV |  |  |  |  | 26.25
107 | BM-TV_24 | Nguyen H. P. | 2002 |  |  | 23.26 |  |  |  |  | TV | TV |  |  |  |  | 28.73
108 | BM-NG_01 | Thieu M. A. | 1991 |  |  |  | 23.74 |  |  |  | NG | NG |  | 23.61
109 | BM-NG_02 | Le T. T. | 2000 |  |  |  | 33.02 |  |  |  | NG | NG |  | 31.72
110 | BM-CTNG_04 | Nguyen X. L. | 1984 |  |  |  | 35.4 | 30.53 |  |  | CT+NG | CT+NG |  | 35.45 | 31.81
111 | BM-CTNG_06 | Nguyen T. K. | 1992 |  |  |  | 26.95 | 22.54 |  |  | CT+NG | CT+NG |  | 26.4 | 24.49
112 | BM-CTNG_08 | Dinh T. D. | 1974 |  |  |  | 39.96 | 30.14 |  |  | CT+NG | CT+NG |  | 36.51 | 32.34
113 | BM-CTNG_10 | Ta V. T. | 1988 |  |  |  | 29.87 | 41.06 |  |  | CT+NG | CT+NG |  | 26.19 | 37.34
114 | BM-CTNG_13 | Dang H. D. | 1984 |  |  |  | 26.43 | 32.73 |  |  | CT+NG | CT+NG |  | 33.91 | 33.56
115 | BM-NG_14 | Vo H. M. | 1990 |  |  |  | 33.11 |  |  |  | NG | NG |  | 37.15
116 | BM-MG_03 | Nguyen T.M. | 1990 | [6 empty] | 31.19 | MG | MG | [5 empty] | 28.33
117 | BM-MG_04 | Dao N. M. | 1976 | [6 empty] | 32.25 | MG | MG | [5 empty] | 29.2
118 | BM-MG_05 | Vu M. D. | 1972 | [6 empty] | 32.82 | MG | MG | [5 empty] | 30.6
119 | BM-MG_06 | Nguyen D. T. | 1988 | [6 empty] | 30.84 | MG | MG | [5 empty] | 29.04
120 | BM-MG_07 | Le Van M. T. | 1988 | [6 empty] | 33.06 | MG | MG | [5 empty] | 31.49
121 | BM-MG_08 | Nguyen T.H | 1985 | [6 empty] | 33.07 | MG | MG | [5 empty] | 31.92
122 | BM-MG_09 | Dang T.V | 1990 | [6 empty] | 33.43 | MG | MG | [5 empty] | 32.04
123 | BM-MG_10 | Pham T.N.A | 1977 | [6 empty] | 32.58 | MG | MG | [5 empty] | 30.1
"non"= negative
STT | Code | Full name | Born | "In house" kit result |  |  | STD FTD 9 |  | Allplex™STI Essential Assay
|  |  |  | HSV-1 (Ct) | HSV-2(Ct) | MH (Ct) | HSV-1 (Ct) | HSV-2(Ct) | MH (Ct)
1 | BM-HSV1- 01 | Bui V.D | 1970 | 32.77 |  |  | P
2 | BM-HSV1- 02 | Nguyen D.T | 1960 | 27.23 |  |  | P
3 | BM-HSV1- 03 | Dinh V.T | 1960 | 29.59 |  |  | P
4 | BM-HSV1- 04 | Dinh V.Th | 1960 | 30.57 |  |  | P
5 | BM-HSV1- 05 | Nguyen D.T | 1960 | 30.16 |  |  | P
6 | BM-HSV1- 06 | Truong B.D | 1972 | 14.53 |  |  | P
7 | BM-HSV1- 07 | Luu V.H | 1974 | 16.63 |  |  | P
8 | BM-HSV1- 08 | Truong B.D | 1972 | 24.64 |  |  | P
9 | BM-HSV1- 09 | Hoang T.T.H | 1992 | 19.29 |  |  | P
10 | BM-HSV1- 10 | Nguyen T.N | 2000 | 17 |  |  | P
11 | BM-HSV1- 11 | Pham V.P | 1949 | 18.47 |  |  | P
12 | BM-HSV1- 12 | Khuong V.T | 1953 | 18.36 |  |  | P
13 | BM-HSV1- 13 | Phan B.M | 1957 | 19.32 |  |  | P
14 | BM-HSV1- 14 | Khuong V.T | 1953 | 19.46 |  |  | P
15 | BM-HSV1- 15 | Dao T.B | 1970 | 17.91 |  |  | P
16 | BM-HSV1- 16 | Doan T.N | 1996 | 18.79 |  |  | P
17 | BM-HSV2-01 | Doan X.D | 1963 |  | 28.62 |  |  | P
18 | BM-HSV2-02 | Tran T.L | 1978 |  | 29.18 |  |  | P
19 | BM-HSV2-03 | Pham V.D | 1964 |  | 33 |  |  | P
20 | BM-HSV2-04 | Trinh Q.M | 1964 |  | 32.58 |  |  | P
21 | BM-HSV2-05 | Mai T.V | 1953 |  | 31.45 |  |  | P
22 | BM-HSV2-06 | Nguyen T.H | 1950 |  | 31.46 |  |  | P
23 | BM-HSV2-07 | Nguyen X.S | 1959 |  | 32.59 |  |  | P
24 | BM-HSV2-08 | Nguyen T.H | 1950 |  | 31.36 |  |  | P
25 | BM-HSV2-09 | Do D.G | 1946 |  | 30.16 |  |  | P
26 | BM-HSV2-10 | Nguyen V.T | 1996 |  | 29.04 |  |  | P
27 | BM-HSV2-11 | Nguyen T.H | 1988 |  | 28.68 |  |  | P
28 | BM-HSV2-12 | Luc T.H | 1974 |  | 27.67 |  |  | P
29 | BM-MH_01 | Bui T.P.A | 1987 |  |  | 27.38 |  |  | P
30 | BM-MH_02 | Vo T.N.H | 1999 |  |  | 23.7 |  |  | P
31 | BM-MH_03 | Vu N. P. | 1991 |  |  | 27.13 |  |  | P
32 | BM-MH_04 | Phung P. D. | 1973 |  |  | 16.2 |  |  | P
33 | BM-MH_05 | D. B. N. | 1985 |  |  | 20.51 |  |  | P
34 | BM-MH_06 | Nguyen T. N. A. | 1993 |  |  | 24.28 |  |  | P
35 | BM-MH_07 | Do M. T. | 1992 |  |  | 31.17 |  |  | P
36 | BM-MH_08 | Mai P. U. V. | 1975 |  |  | 19.17 |  |  | P
37 | BM-MH_09 | Trinh D. H. | 1979 |  |  | 20.16 |  |  | P
38 | BM-MH_10 | Hoang T. L. | 1978 |  |  | 16.46 |  |  | P
39 | BM-MH_11 | Nguyen T. A. | 1993 |  |  | 19.73 |  |  | P
40 | BM-MH_12 | Nguyen A. N. | 1992 |  |  | 30.36 |  |  | P
41 | BM-MH_13 | Phi B. L. | 1995 |  |  | 28.6 |  |  | P
42 | BM-MH_14 | Luu B. Q. | 2000 |  |  | 24.2 |  |  | P
43 | BM-MH_15 | Nguyen T.T.N | 2001 |  |  | 23.08 |  |  | P
44 | BM_AT_01 | Nguyen T.H | 1983 | NA | NA | NA | NA | NA | NA
45 | BM_AT_02 | Nguyen H.M | 1962 | NA | NA | NA | NA | NA | NA
46 | BM_AT_03 | Nguyen T.X | 1991 | NA | NA | NA | NA | NA | NA
47 | BM_AT_04 | Pham H.T | 1991 | NA | NA | NA | NA | NA | NA
48 | BM_AT_05 | Vu T.T.N | 1983 | NA | NA | NA | NA | NA | NA
49 | BM_AT_06 | Nguyen T.L | 1974 | NA | NA | NA | NA | NA | NA
50 | BM_AT_07 | Tu T.H | 1980 | NA | NA | NA | NA | NA | NA
51 | BM_AT_08 | Cao T.T | 1983 | NA | NA | NA | NA | NA | NA
52 | BM_AT_09 | Vu T.H | 1970 | NA | NA | NA | NA | NA | NA
53 | BM_AT_10 | Lai T.T | 1973 | NA | NA | NA | NA | NA | NA
54 | BM_AT_11 | Nguyen T.T.T | 1988 | NA | NA | NA | NA | NA | NA
55 | BM_AT_12 | Pham T.H | 1975 | NA | NA | NA | NA | NA | NA
56 | BM_AT_13 | Nguyen T.T.L | 1986 | NA | NA | NA | NA | NA | NA
57 | BM_AT_14 | Vu T.T.H | 1981 | NA | NA | NA | NA | NA | NA
58 | BM_AT_15 | Vu T.C | 1972 | NA | NA | NA | NA | NA | NA
59 | BM_AT_16 | Nguyen T.B.T | 1976 | NA | NA | NA | NA | NA | NA
60 | BM_AT_17 | Nguyen T.K.H | 1984 | NA | NA | NA | NA | NA | NA
61 | BM_AT_18 | Nguyen T.M | 1962 | NA | NA | NA | NA | NA | NA
62 | BM_AT_19 | Do T.L.P | 1995 | NA | NA | NA | NA | NA | NA
63 | BM_AT_20 | Nguyen P.L | 1990 | NA | NA | NA | NA | NA | NA
64 | BM_AT_21 | Pham T.T | 1976 | NA | NA | NA | NA | NA | NA
65 | BM_AT_22 | Nguyen H.L | 1993 | NA | NA | NA | NA | NA | NA
66 | BM_AT_23 | Nguyen T.M | 1988 | NA | NA | NA | NA | NA | NA
67 | BM_AT_24 | Do T.H | 1982 | NA | NA | NA | NA | NA | NA
68 | BM_AT_25 | Lo H.A | 1997 | NA | NA | NA | NA | NA | NA
69 | BM_AT_26 | Vu H.H | 1975 | NA | NA | NA | NA | NA | NA
70 | BM_AT_27 | Pham T.V | 1974 | NA | NA | NA | NA | NA | NA
71 | BM_AT_28 | Luong T.N | 1994 | NA | NA | NA | NA | NA | NA
72 | BM_AT_29 | Nguyen T.T | 1984 | NA | NA | NA | NA | NA | NA
73 | BM_AT_30 | Nguyen T.H.T | 1983 | NA | NA | NA | NA | NA | NA
74 | BM_AT_31 | Ngo T.Q | 1978 | NA | NA | NA | NA | NA | NA
75 | BM_AT_32 | Pham T.H | 1984 | NA | NA | NA | NA | NA | NA
76 | BM_AT_33 | Nguyen T.T.T | 1990 | NA | NA | NA | NA | NA | NA
77 | BM_AT_34 | Nguyen T.D | 1977 | NA | NA | NA | NA | NA | NA
78 | BM_AT_35 | Nham T.H.H | 1978 | NA | NA | NA | NA | NA | NA
79 | BM_AT_36 | Ngo T.T | 1980 | NA | NA | NA | NA | NA | NA
80 | BM_AT_37 | Le T.G | 1993 | NA | NA | NA | NA | NA | NA
81 | BM_AT_38 | Nguyen T.K.D | 1970 | NA | NA | NA | NA | NA | NA
82 | BM_AT_39 | Do H.N | 1995 | NA | NA | NA | NA | NA | NA
83 | BM_AT_40 | Dang T.T.H | 1990 | NA | NA | NA | NA | NA | NA
84 | BM_AT_41 | Nguyen T.H | 1991 | NA | NA | NA | NA | NA | NA
85 | BM_AT_42 | Dinh T.H.N | 1993 | NA | NA | NA | NA | NA | NA
86 | BM_AT_43 | Truong M.N | 1988 | NA | NA | NA | NA | NA | NA
87 | BM_AT_44 | Nguyen T.T | 1953 | NA | NA | NA | NA | NA | NA
88 | BM_AT_45 | Do T.H | 1993 | NA | NA | NA | NA | NA | NA
89 | BM_AT_46 | Nguyen D.M | 1998 | NA | NA | NA | NA | NA | NA
90 | BM_AT_47 | Trinh T.H.L | 1996 | NA | NA | NA | NA | NA | NA
91 | BM_AT_48 | Tran T.M | 1964 | NA | NA | NA | NA | NA | NA
92 | BM_AT_49 | Tran T.P | 1975 | NA | NA | NA | NA | NA | NA
93 | BM_AT_50 | Nguyen T.V | 1985 | NA | NA | NA | NA | NA | NA
94 | BM_AT_51 | Nghiem T.N | 1969 | NA | NA | NA | NA | NA | NA
95 | BM_AT_52 | Mach T.V | 1990 | NA | NA | NA | NA | NA | NA
96 | BM_AT_53 | Nguyen T.H | 1998 | NA | NA | NA | NA | NA | NA
97 | BM_AT_54 | Nguyen T.N | 1994 | NA | NA | NA | NA | NA | NA
98 | BM_AT_55 | Nguyen T.L | 1991 | NA | NA | NA | NA | NA | NA
99 | BM_AT_56 | Chu T.L | 1987 | NA | NA | NA | NA | NA | NA
100 | BM_AT_57 | Phung T.H | 1984 | NA | NA | NA | NA | NA | NA
101 | BM_AT_58 | Ngo T.T.C | 1990 | NA | NA | NA | NA | NA | NA
102 | BM_AT_59 | Nguyen T.P.T | 1998 | NA | NA | NA | NA | NA | NA
103 | BM_AT_60 | Vu T.L | 1978 | NA | NA | NA | NA | NA | NA
104 | BM_AT_61 | Nguyen H.H | 1993 | NA | NA | NA | NA | NA | NA
|  | "NA" | Not available
|  | "P" | Positive sample

## Table 4
Table 4. Limit of detection of “in-house” multiplex real-time PCR assay | [21 empty] | Pathogen name | No. of replicates (%) for positive standards of real-time PCR |  |  |  |  | LOD, copies/reactions
Target pathogen | No. of replicates (%) for standard samples, copies/reaction | [26 empty] | (95% CI)
| A | B | C | D | E | [17 empty] | 500 copies | 50 copies | 10 copies | 5 copies | 1 copy
| 500 | 50 | 10 | 5 | 1 | Total | [15 empty] | G. vaginalis | 3/3 (100%) | 20/20 (100%) | 19/20 (95%) | 9/20 (45%) | 2/20 (10%) | 17.1 (11.69-25.15)
GV | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | N. gonorrhoeae | 3/3 (100%) | 20/20 (100%) | 20/20 (100%) | 7/20 (35%) | 1/20 (5%) | 18.6 (12.62-27.32)
NG | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | C. trachomatis | 3/3 (100%) | 20/20 (100%) | 10/20 (50%) | 8/20 (40%) | 0/20 (0%) | 37.2 (24.86-55.81)
TV | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | HSV-1 | 3/3 (100%) | 20/20 (100%) | 13/20 (65%) | 12/20 (60%) | 2/20 (10%) | 22.8 (15.36-33.84)
HSV-1 | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | C. albicans | 3/3 (100%) | 20/20 (100%) | 20/20 (100%) | 10/20 (50%) | 1/20 (5%) | 15.5 (10.58-22.60)
CA | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | T. vaginalis | 3/3 (100%) | 20/20 (100%) | 20/20 (100%) | 17/20 (85%) | 4/20 (20%) | 8.2 (5.56-12.19)
CT | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | M. hominis | 3/3 (100%) | 20/20 (100%) | 17/20 (85%) | 7/20 (15%) | 0/20 (0%) | 24.5 (16.49-36.53)
MH | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | HSV-2 | 3/3 (100%) | 20/20 (100%) | 10/20 (50%) | 3/20 (15%) | 0/20 (0%) | 48.3 (32.27-72.40)
HSV-2 | 3 | 20 | 20 | 20 | 20 | 83 | [15 empty] | M. genitalium | 3/3 (100%) | 20/20 (100%) | 7/20 (35%) | 3/20 (15%) | 0/20 (0%) | 57.8 (38.56-86.75)
MG | 3 | 20 | 20 | 20 | 20 | 83
Total | [5 empty] | 747
G. vaginalis (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 32.46 | 33.04 | 32.7 | [17 empty] | 3/3
50 copies | 35.91 | 36.5 | 35.26 | 36.65 | 37.02 | 34.46 | 35 | 36.92 | 36.48 | 36.37 | 37.15 | 37.41 | 37.03 | 36.96 | 36.19 | 37.75 | 36.34 | 34.88 | 36.04 | 36.75 | 20/20
10 copies | 36.92 | 36.8 | 36.79 | 37.2 | NA | 36.12 | 36.61 | 38.21 | 37.22 | 37.96 | 37.09 | 38.37 | 36.33 | 35.84 | 35.48 | 36.24 | 36 | 36.49 | 38.35 | 35.8 | 19/20
5 copies | 41.31 | NA | NA | NA | 39.54 | NA | 37.56 | NA | 37.25 | 38.6 | 36.94 | NA | NA | 37.64 | NA | 39.16 | 37.34 | NA | NA | NA | 9/20
1 copy | NA | NA | NA | NA | NA | 40.77 | NA | NA | NA | NA | NA | NA | NA | NA | 40.18 | NA | NA | NA | NA | NA | 2/20
N. gonorrhoeae (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 34.23 | 33.73 | 34.9 | [17 empty] | 3/3
50 copies | 36.02 | 37.93 | 37.38 | 37.15 | 37.61 | 38.1 | 41.02 | 36.11 | 40.79 | 38.21 | 36.2 | 37.9 | 37.67 | 39.58 | 36.37 | 37.89 | 37.82 | 37.66 | 37.25 | 36.92 | 20/20
10 copies | 37.37 | 38.52 | 37.56 | 38.02 | 38.5 | 36.53 | 38.06 | 41.21 | 40.76 | 38.2 | 36.75 | 36.58 | 37.45 | 37.07 | 38.07 | 36.96 | 36.44 | 39.06 | 38.91 | 37.67 | 20/20
5 copies | NA | 40.85 |  |  | 41.1 | 39.76 |  |  |  |  | 38.2 |  |  |  | 39.98 |  |  | 38.8 | 37.89 |  | 7/20
1 copy | 39.64 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1/20
C. trachomatis(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 34.11 | 33.32 | 32.3
50 copies | 36.6 | 35.43 | 35.47 | 34.98 | 35 | 34.81 | 37.22 | 35.77 | 36 | 36.06 | 35.1 | 34.85 | 34.92 | 35.47 | 34.56 | 34.86 | 34.72 | 36.02 | 36.05 | 34.45 | 20/20
10 copies | NA | 36.69 | 35.69 | 35.7 | NA | NA | 37 | NA | 35.75 | NA | 37.03 | NA | NA | NA | 37.9 | 35.86 | 38 | NA | 36.33 | NA | 10/20
5 copies | NA | NA | NA | NA | 37.86 | 36.6 | 36.23 | NA | 37.71 | 37.06 | NA | 36.6 | NA | 39 | NA | 39.75 | NA | NA | NA | NA | 8/20
1 copy | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20
HSV-1(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 32.96 | 33.47 | 32.66 | [17 empty] | 3/3
50 copies | 36.62 | 37.46 | 36.56 | 37.09 | 36.2 | 37.72 | 36.79 | 37.2 | 37.5 | 36.34 | 37.21 | 35.98 | 37.7 | 36.76 | 37.91 | 38.08 | 35.52 | 38.02 | 35.67 | 35.49 | 20/20
10 copies | 41.65 | 37.41 | NA | 38.07 | 40.46 | NA | 38.73 | 41.02 | 40.23 | 38.96 | NA | NA | NA | 37.89 | 37.95 | NA | 40.05 | NA | 38.94 | 37.03 | 13/20
5 copies | NA | 38.99 | 40.24 | 39.34 | 39.02 | 39.1 | 38.62 | 37.8 | 38.87 | 37.87 | NA | 37.5 | NA | NA | NA | NA | NA | NA | 38.77 | 38.06 | 12/20
1 copy | 38.3 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 40.6 | NA | NA | NA | NA | NA | 2/20
HSV-2(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.42 | 33.82 | 33.36 | [17 empty] | 3/3
50 copies | 36.71 | 35.11 | 34.51 | 35.75 | 35.41 | 34.94 | 36.14 | 34.75 | 35.48 | 35.66 | 34.35 | 35.19 | 34.56 | 35.77 | 34.63 | 36.07 | 35.96 | 34.68 | 34.82 | 35.02 | 20/20
10 copies | NA | NA | NA | 35.69 | 35.73 | NA | 35.32 | NA | 35.87 | 36.15 | 36.11 | 35.28 | 36.76 | 36.31 | NA | 38.03 | NA | NA | NA | NA | 10/20
5 copies | NA | NA | NA | NA | NA | NA | NA | NA | NA | 37.71 | NA | NA | NA | 35.9 | NA | NA | NA | 35.47 | NA | NA | 3/20
1 copy | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20
C. albicans(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.87 | 33.11 | 33.17 | [17 empty] | 3/3
50 copies | 37.26 | 37.33 | 37.43 | 37.65 | 37.35 | 36.16 | 37.49 | 36.91 | 37.67 | 37.52 | 37.17 | 37.86 | 37.71 | 37.66 | 36.28 | 38.68 | 36.21 | 35.84 | 37.87 | 38.2 | 20/20
10 copies | NA | NA | 38.72 | NA | NA | 38.87 | 40.71 | NA | 40.81 | 39.54 | NA | NA | 39.68 | 40.28 | 39.09 | 37.69 | 39.81 | NA | NA | NA | 10/20
5 copies | NA | NA | NA | 39.79 | NA | 38.29 | 39.19 | NA | NA | NA | NA | 36.91 | NA | 38 | 37.15 | NA | 37.65 | 39.11 | 38.34 | 38.77 | 10/20
1 copy | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 40.32 | 1/20
T. vaginalis (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.57 | 33.24 | 32.23 | [17 empty] | 3/3
50 copies | 37.77 | 36.23 | 36.07 | 35.46 | 35.36 | 35.66 | 36.76 | 35.52 | 36.2 | 36.51 | 35.76 | 36.45 | 36.86 | 35 | 36.06 | 36.27 | 36.28 | 35.43 | 35.13 | 36.98 | 20/20
10 copies | 39.12 | 38.17 | 37.76 | 38.16 | 39.08 | 37.15 | 38.16 | NA | 37.4 | NA | 36.82 | 37.87 | 36.9 | 36.63 | 37.94 | 38.59 | 37.75 | NA | 38.03 | 38.85 | 17/20
5 copies | 39.31 | 38.35 | NA | NA | NA | 39 | 38.13 | 38.29 | 38.53 | 39 | 38.37 | 38.75 | 37.01 | 39.75 | 38.6 | 37.99 | 38.92 | 38.47 | 38.61 | 37.39 | 17/20
1 copy | 41.87 | NA | NA | NA | NA | NA | NA | 38.57 | NA | NA | 37.55 | NA | NA | NA | NA | 38.65 | NA | NA | NA | NA | 4/20
M. hominis (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.61 | 33.7 | 33.68 | [17 empty] | 3/3
50 copies | 36.66 | 36.51 | 35.94 | 36.32 | 36.26 | 36.9 | 35.89 | 35.7 | 36.1 | 35.46 | 36.27 | 35.88 | 36.24 | 35.86 | 35.78 | 36.01 | 36.05 | 36.34 | 35.46 | 36.31 | 20/20
10 copies | 36.86 | 37.1 | 38.29 | 37.95 | 37.54 | NA | 37.42 | 39.74 | 39.65 | 37.13 | 37.5 | NA | 38.3 | 40.24 | 38.6 | 38.65 | 37.43 | 41.19 | 38.95 | NA | 17/20
5 copies | NA | 41.75 | 36.85 | 40.57 | 37.83 | NA | NA | NA | NA | NA | NA | 39.86 | 39.93 | 40.02 | NA | NA | NA | NA | NA | NA | 7/20
1 copy | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20
M. genitalium (Ct)
Copies | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 34.07 | 34.64 | 35.2 | [17 empty] | 3/3
50 copies | 35.87 | 34.85 | 35.54 | 34.24 | 36.86 | 36.18 | 33.88 | 35.57 | 34.71 | 34.82 | 34.87 | 34.94 | 35.17 | 34.16 | 34.38 | 34.86 | 34.7 | 35.38 | 34.78 | 36 | 20/20
10 copies | 37.79 | 35.77 | 37.07 | NA | NA | 36.88 | NA | NA | NA | NA | NA | NA | 36.08 | 35.68 | 38.44 | NA | NA | NA | NA | NA | 7/20
5 copies | 41.83 | NA | NA | NA | NA | NA | NA | NA | NA | 38.5 | NA | NA | NA | NA | NA | 36.18 | NA | NA | NA | NA | 3/20
1 copy | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20
"NA": not available

## Table 5-detailed
Enterprise reference panel |  |  | Repeatability test result | [26 empty] | Reproducibility test result
|  |  | Batch 1 | [8 empty] | Batch 2 | [8 empty] | Batch 3
|  |  | Replicate | [5 empty] | Mean1 | SD1 | CV1 value (%) | Replicate | [5 empty] | Mean2 | SD2 | CV2 value (%) | Replicate | [5 empty] | Mean3 | SD3 | CV3 value (%) | Batch 1 | Batch 2 | Batch 3 | Mean | SD | CV value (%)
|  |  | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | Mean 1 | Mean 2 | Mean 2
Nine STD pathogens of enterprise reference panel | GV | 50 copies | 35.91 | 36.5 | 36.5 | 36.65 | 37.02 | 34.46 | 36.1733 | 0.9124 | 2.5223 | 35 | 36.92 | 36.48 | 36.37 | 37.15 | 37.41 | 36.555 | 0.85746 | 2.3457 | 37.03 | 36.96 | 36.19 | 37.75 | 36.34 | 34.88 | 36.525 | 0.97953 | 2.6818 | 36.1733 | 36.555 | 36.525 | 36.4178 | 0.85558 | 2.3493
| NG | 50 copies | 36.02 | 37.93 | 37.38 | 37.15 | 37.61 | 38.1 | 37.365 | 0.74498 | 1.9938 | 37.25 | 36.11 | 36.92 | 38.21 | 36.2 | 37.9 | 37.0983 | 0.86222 | 2.3241 | 37.67 | 39.58 | 36.37 | 37.89 | 37.82 | 37.66 | 37.8317 | 1.0249 | 2.7092 | 37.365 | 37.0983 | 37.8317 | 37.4317 | 0.86274 | 2.3048
| CT | 50 copies | 36.6 | 35.43 | 35.47 | 34.98 | 35 | 34.81 | 35.3817 | 0.65242 | 1.844 | 37.22 | 35.77 | 36 | 36.06 | 35.1 | 34.85 | 35.8333 | 0.83794 | 2.3384 | 34.92 | 35.47 | 34.56 | 34.86 | 34.72 | 36.02 | 35.0917 | 0.54934 | 1.5654 | 35.3817 | 35.8333 | 35.0917 | 35.4356 | 0.70017 | 1.9759
| HSV1 | 50 copies | 36.62 | 37.46 | 36.56 | 37.09 | 36.2 | 37.72 | 36.9417 | 0.58235 | 1.5764 | 36.79 | 37.2 | 37.5 | 36.34 | 37.21 | 35.98 | 36.8367 | 0.58209 | 1.5802 | 37.7 | 36.76 | 37.91 | 38.08 | 35.52 | 38.02 | 37.3317 | 1.0112 | 2.7088 | 36.9417 | 36.8367 | 37.3317 | 37.0367 | 0.71953 | 1.9427
| CA | 50 copies | 37.26 | 37.33 | 37.43 | 37.65 | 37.35 | 36.16 | 37.1967 | 0.52534 | 1.4123 | 37.49 | 36.91 | 37.67 | 37.52 | 37.17 | 37.86 | 37.4367 | 0.34407 | 0.91908 | 37.71 | 37.66 | 36.28 | 38.68 | 36.21 | 35.84 | 37.0633 | 1.1159 | 3.0108 | 37.1967 | 37.4367 | 37.0633 | 37.2322 | 0.69232 | 1.8595
| TV | 50 copies | 37.77 | 36.23 | 36.07 | 35.46 | 35.36 | 35.66 | 36.0917 | 0.88962 | 2.4649 | 36.76 | 35.52 | 36.2 | 36.51 | 35.76 | 36.45 | 36.2 | 0.47502 | 1.3122 | 36.86 | 35 | 36.06 | 36.27 | 36.28 | 35.43 | 35.9833 | 0.66599 | 1.8508 | 36.0917 | 36.2 | 35.9833 | 36.0917 | 0.64308 | 1.7818
| MH | 50 copies | 36.66 | 36.51 | 35.94 | 36.32 | 36.26 | 36.9 | 36.4317 | 0.33505 | 0.91966 | 35.89 | 35.7 | 36.1 | 35.46 | 36.27 | 35.88 | 35.8833 | 0.28612 | 0.79737 | 36.24 | 35.86 | 35.78 | 36.01 | 36.05 | 36.34 | 36.0467 | 0.21482 | 0.59594 | 36.4317 | 35.8833 | 36.0467 | 36.1206 | 0.34581 | 0.95737
| HSV2 | 50 copies | 36.71 | 35.11 | 34.51 | 35.75 | 35.41 | 34.94 | 35.405 | 0.76503 | 2.1608 | 36.14 | 34.75 | 35.48 | 35.66 | 34.35 | 35.19 | 35.2617 | 0.64434 | 1.8273 | 34.56 | 35.77 | 34.63 | 36.07 | 35.96 | 34.68 | 35.2783 | 0.72491 | 2.0548 | 35.405 | 35.2617 | 35.2783 | 35.315 | 0.65419 | 1.8525
| MG | 50 copies | 35.87 | 34.85 | 35.54 | 34.24 | 36.86 | 36.18 | 35.59 | 0.93936 | 2.6394 | 33.88 | 35.57 | 34.71 | 34.82 | 34.87 | 34.94 | 34.7983 | 0.54264 | 1.5594 | 35.17 | 34.16 | 34.38 | 34.86 | 34.7 | 35.38 | 34.775 | 0.46242 | 1.3297 | 35.59 | 34.7983 | 34.775 | 35.0544 | 0.72788 | 2.0764
G. vaginalis (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 32.46 | 33.04 | Replicate 1: 32.7 | [12 empty] | 3/3
50 copies | 35.91 | 36.5 | 36.5 | 36.65 | 37.02 | 34.46 | 35 | 36.92 | 36.48 | 36.37 | 37.15 | 37.41 | 37.03 | 36.96 | 36.19 | 37.75 | 36.34 | 34.88 | 36.04 | 36.75 | 20/20
10 copies | 36.92 | 36.8 | 36.79 | 37.2 | NA | 36.12 | 36.61 | 38.21 | 37.22 | 37.96 | 37.09 | 38.37 | 36.33 | 35.84 | 35.48 | 36.24 | 36 | 36.49 | 38.35 | 35.8 | 19/20
5 copies | 41.31 | NA | NA | NA | 39.54 | NA | 37.56 | NA | 37.25 | 38.6 | 36.94 | NA | NA | 37.64 | NA | 39.16 | 37.34 | NA | NA | NA | 9/20
1 copy | NA | NA | NA | NA | NA | 40.77 | NA | NA | NA | NA | NA | NA | NA | NA | 40.18 | NA | NA | NA | NA | NA | 2/20
N. gonorrhoeae (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 34.23 | 33.73 | Replicate 1: 34.9 | [12 empty] | 3/3
50 copies | 36.02 | 37.93 | 37.38 | 37.15 | 37.61 | 38.1 | 41.02 | 36.11 | 40.79 | 38.21 | 36.2 | 37.9 | 37.67 | 39.58 | 36.37 | 37.89 | 37.82 | 37.66 | 37.25 | 36.92 | 20/20
10 copies | 37.37 | 38.52 | 37.56 | 38.02 | 38.5 | 36.53 | 38.06 | 41.21 | 40.76 | 38.2 | 36.75 | 36.58 | 37.45 | 37.07 | 38.07 | 36.96 | 36.44 | 39.06 | 38.91 | 37.67 | 20/20
5 copies | NA | 40.85 | Replicate 3: 41.1; Replicate 4: 39.76 |  |  | 38.2 |  |  |  | 39.98 |  |  | 38.8 | 37.89 |  | 7/20
1 copy | 39.64 | NA | [6 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1/20 | bar: Replicate 1 to Replicate 6
C. trachomatis(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 34.11 | 33.32 | Replicate 1: 32.3
50 copies | 36.6 | 35.43 | 35.47 | 34.98 | 35 | 34.81 | 37.22 | 35.77 | 36 | 36.06 | 35.1 | 34.85 | 34.92 | 35.47 | 34.56 | 34.86 | 34.72 | 36.02 | 36.05 | 34.45 | 20/20
10 copies | NA | 36.69 | 35.69 | 35.7 | NA | NA | 37 | NA | 35.75 | NA | 37.03 | NA | NA | NA | 37.9 | 35.86 | 38 | NA | 36.33 | NA | 10/20
5 copies | NA | NA | NA | NA | 37.86 | 36.6 | 36.23 | NA | 37.71 | 37.06 | NA | 36.6 | NA | 39 | NA | 39.75 | NA | NA | NA | NA | 8/20
1 copy | NA | NA | [6 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20 | bar: Replicate 1 to Replicate 6
HSV-1(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 32.96 | 33.47 | Replicate 1: 32.66 | [12 empty] | 3/3
50 copies | 36.62 | 37.46 | 36.56 | 37.09 | 36.2 | 37.72 | 36.79 | 37.2 | 37.5 | 36.34 | 37.21 | 35.98 | 37.7 | 36.76 | 37.91 | 38.08 | 35.52 | 38.02 | 35.67 | 35.49 | 20/20
10 copies | 41.65 | 37.41 | NA | 38.07 | 40.46 | NA | 38.73 | 41.02 | 40.23 | 38.96 | NA | NA | NA | 37.89 | 37.95 | NA | 40.05 | NA | 38.94 | 37.03 | 13/20
5 copies | NA | 38.99 | 40.24 | 39.34 | 39.02 | 39.1 | 38.62 | 37.8 | 38.87 | 37.87 | NA | 37.5 | NA | NA | NA | NA | NA | NA | 38.77 | 38.06 | 12/20
1 copy | 38.3 | NA | [6 empty] | NA | NA | NA | NA | NA | NA | 40.6 | NA | NA | NA | NA | NA | 2/20 | bar: Replicate 1 to Replicate 6
HSV-2(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.42 | 33.82 | Replicate 1: 33.36 | [12 empty] | 3/3
50 copies | 36.71 | 35.11 | 34.51 | 35.75 | 35.41 | 34.94 | 36.14 | 34.75 | 35.48 | 35.66 | 34.35 | 35.19 | 34.56 | 35.77 | 34.63 | 36.07 | 35.96 | 34.68 | 34.82 | 35.02 | 20/20
10 copies | NA | NA | NA | 35.69 | 35.73 | NA | 35.32 | NA | 35.87 | 36.15 | 36.11 | 35.28 | 36.76 | 36.31 | NA | 38.03 | NA | NA | NA | NA | 10/20
5 copies | NA | NA | [6 empty] | NA | 37.71 | NA | NA | NA | 35.9 | NA | NA | NA | 35.47 | NA | NA | 3/20 | bar: Replicate 1 to Replicate 6
1 copy | NA | NA | [6 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20 | bar: Replicate 1 to Replicate 6
C. albicans(Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.87 | 33.11 | Replicate 1: 33.17 | [12 empty] | 3/3
50 copies | 37.26 | 37.33 | 37.43 | 37.65 | 37.35 | 36.16 | 37.49 | 36.91 | 37.67 | 37.52 | 37.17 | 37.86 | 37.71 | 37.66 | 36.28 | 38.68 | 36.21 | 35.84 | 37.87 | 38.2 | 20/20
10 copies | NA | NA | 38.72 | NA | NA | 38.87 | 40.71 | NA | 40.81 | 39.54 | NA | NA | 39.68 | 40.28 | 39.09 | 37.69 | 39.81 | NA | NA | NA | 10/20
5 copies | NA | NA | NA | 39.79 | NA | 38.29 | 39.19 | NA | NA | NA | NA | 36.91 | NA | 38 | 37.15 | NA | 37.65 | 39.11 | 38.34 | 38.77 | 10/20
1 copy | NA | NA | [6 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 40.32 | 1/20 | bar: Replicate 1 to Replicate 6
T. vaginalis (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.57 | 33.24 | Replicate 1: 32.23 | [12 empty] | 3/3
50 copies | 37.77 | 36.23 | 36.07 | 35.46 | 35.36 | 35.66 | 36.76 | 35.52 | 36.2 | 36.51 | 35.76 | 36.45 | 36.86 | 35 | 36.06 | 36.27 | 36.28 | 35.43 | 35.13 | 36.98 | 20/20
10 copies | 39.12 | 38.17 | 37.76 | 38.16 | 39.08 | 37.15 | 38.16 | NA | 37.4 | NA | 36.82 | 37.87 | 36.9 | 36.63 | 37.94 | 38.59 | 37.75 | NA | 38.03 | 38.85 | 17/20
5 copies | 39.31 | 38.35 | NA | NA | NA | 39 | 38.13 | 38.29 | 38.53 | 39 | 38.37 | 38.75 | 37.01 | 39.75 | 38.6 | 37.99 | 38.92 | 38.47 | 38.61 | 37.39 | 17/20
1 copy | 41.87 | NA | NA | NA | NA | NA | NA | 38.57 | NA | NA | 37.55 | NA | NA | NA | NA | 38.65 | NA | NA | NA | NA | 4/20
M. hominis (Ct)
Time | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 33.61 | 33.7 | Replicate 1: 33.68 | [12 empty] | 3/3
50 copies | 36.66 | 36.51 | 35.94 | 36.32 | 36.26 | 36.9 | 35.89 | 35.7 | 36.1 | 35.46 | 36.27 | 35.88 | 36.24 | 35.86 | 35.78 | 36.01 | 36.05 | 36.34 | 35.46 | 36.31 | 20/20
10 copies | 36.86 | 37.1 | 38.29 | 37.95 | 37.54 | NA | 37.42 | 39.74 | 39.65 | 37.13 | 37.5 | NA | 38.3 | 40.24 | 38.6 | 38.65 | 37.43 | 41.19 | 38.95 | NA | 17/20
5 copies | NA | 41.75 | 36.85 | 40.57 | 37.83 | NA | NA | NA | NA | NA | NA | 39.86 | 39.93 | 40.02 | NA | NA | NA | NA | NA | NA | 7/20
1 copy | NA | NA | [6 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20 | bar: Replicate 1 to Replicate 6
M. genitalium (Ct)
Copies | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | Total
500 copies | 34.07 | 34.64 | Replicate 1: 35.2 | [12 empty] | 3/3
50 copies | 35.87 | 34.85 | 35.54 | 34.24 | 36.86 | 36.18 | 33.88 | 35.57 | 34.71 | 34.82 | 34.87 | 34.94 | 35.17 | 34.16 | 34.38 | 34.86 | 34.7 | 35.38 | 34.78 | 36 | 20/20
10 copies | 37.79 | 35.77 | 37.07 | NA | NA | 36.88 | NA | NA | NA | NA | NA | NA | 36.08 | 35.68 | 38.44 | NA | NA | NA | NA | NA | 7/20
5 copies | 41.83 | NA | [6 empty] | NA | 38.5 | NA | NA | NA | NA | NA | 36.18 | NA | NA | NA | NA | 3/20 | bar: Replicate 1 to Replicate 6
1 copy | NA | NA | [6 empty] | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0/20 | bar: Replicate 1 to Replicate 6
"NA": not available

## Table 5-summary
Nine STD pathogens of reference panel |  |  | Repeatability test result | [26 empty] | Reproducibility test result
|  |  | Batch 1 | [8 empty] | Batch 2 | [8 empty] | Batch 3
|  |  | Replicate | [5 empty] | Mean1 | SD1 | CV1 value (%) | Replicate | [5 empty] | Mean2 | SD2 | CV2 value (%) | Replicate | [5 empty] | Mean3 | SD3 | CV3 value (%) | Batch 1 | Batch 2 | Batch 3 | Mean | SD | CV value (%)
|  |  | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 |  |  |  | Mean 1 | Mean 2 | Mean 2
| GV | 50 copies | 35.91 | 36.5 | 36.5 | 36.65 | 37.02 | 34.46 | 36.1733 | 0.9124 | 2.5223 | 35 | 36.92 | 36.48 | 36.37 | 37.15 | 37.41 | 36.555 | 0.85746 | 2.3457 | 37.03 | 36.96 | 36.19 | 37.75 | 36.34 | 34.88 | 36.525 | 0.97953 | 2.6818 | 36.1733 | 36.555 | 36.525 | 36.4178 | 0.21223 | 0.58275
| NG | 50 copies | 36.02 | 37.93 | 37.38 | 37.15 | 37.61 | 38.1 | 37.365 | 0.74498 | 1.9938 | 37.25 | 36.11 | 36.92 | 38.21 | 36.2 | 37.9 | 37.0983 | 0.86222 | 2.3241 | 37.67 | 39.58 | 36.37 | 37.89 | 37.82 | 37.66 | 37.8317 | 1.0249 | 2.7092 | 37.365 | 37.0983 | 37.8317 | 37.4317 | 0.37118 | 0.99163
| CT | 50 copies | 36.6 | 35.43 | 35.47 | 34.98 | 35 | 34.81 | 35.3817 | 0.65242 | 1.844 | 37.22 | 35.77 | 36 | 36.06 | 35.1 | 34.85 | 35.8333 | 0.83794 | 2.3384 | 34.92 | 35.47 | 34.56 | 34.86 | 34.72 | 36.02 | 35.0917 | 0.54934 | 1.5654 | 35.3817 | 35.8333 | 35.0917 | 35.4356 | 0.37376 | 1.0548
| HSV1 | 50 copies | 36.62 | 37.46 | 36.56 | 37.09 | 36.2 | 37.72 | 36.9417 | 0.58235 | 1.5764 | 36.79 | 37.2 | 37.5 | 36.34 | 37.21 | 35.98 | 36.8367 | 0.58209 | 1.5802 | 37.7 | 36.76 | 37.91 | 38.08 | 35.52 | 38.02 | 37.3317 | 1.0112 | 2.7088 | 36.9417 | 36.8367 | 37.3317 | 37.0367 | 0.26082 | 0.70421
| CA | 50 copies | 37.26 | 37.33 | 37.43 | 37.65 | 37.35 | 36.16 | 37.1967 | 0.52534 | 1.4123 | 37.49 | 36.91 | 37.67 | 37.52 | 37.17 | 37.86 | 37.4367 | 0.34407 | 0.91908 | 37.71 | 37.66 | 36.28 | 38.68 | 36.21 | 35.84 | 37.0633 | 1.1159 | 3.0108 | 37.1967 | 37.4367 | 37.0633 | 37.2322 | 0.18919 | 0.50813
| TV | 50 copies | 37.77 | 36.23 | 36.07 | 35.46 | 35.36 | 35.66 | 36.0917 | 0.88962 | 2.4649 | 36.76 | 35.52 | 36.2 | 36.51 | 35.76 | 36.45 | 36.2 | 0.47502 | 1.3122 | 36.86 | 35 | 36.06 | 36.27 | 36.28 | 35.43 | 35.9833 | 0.66599 | 1.8508 | 36.0917 | 36.2 | 35.9833 | 36.0917 | 0.10833 | 0.30016
| MH | 50 copies | 36.66 | 36.51 | 35.94 | 36.32 | 36.26 | 36.9 | 36.4317 | 0.33505 | 0.91966 | 35.89 | 35.7 | 36.1 | 35.46 | 36.27 | 35.88 | 35.8833 | 0.28612 | 0.79737 | 36.24 | 35.86 | 35.78 | 36.01 | 36.05 | 36.34 | 36.0467 | 0.21482 | 0.59594 | 36.4317 | 35.8833 | 36.0467 | 36.1206 | 0.28154 | 0.77943
| HSV2 | 50 copies | 36.71 | 35.11 | 34.51 | 35.75 | 35.41 | 34.94 | 35.405 | 0.76503 | 2.1608 | 36.14 | 34.75 | 35.48 | 35.66 | 34.35 | 35.19 | 35.2617 | 0.64434 | 1.8273 | 34.56 | 35.77 | 34.63 | 36.07 | 35.96 | 34.68 | 35.2783 | 0.72491 | 2.0548 | 35.405 | 35.2617 | 35.2783 | 35.315 | 0.078387 | 0.22196
| MG | 50 copies | 35.87 | 34.85 | 35.54 | 34.24 | 36.86 | 36.18 | 35.59 | 0.93936 | 2.6394 | 33.88 | 35.57 | 34.71 | 34.82 | 34.87 | 34.94 | 34.7983 | 0.54264 | 1.5594 | 35.17 | 34.16 | 34.38 | 34.86 | 34.7 | 35.38 | 34.775 | 0.46242 | 1.3297 | 35.59 | 34.7983 | 34.775 | 35.0544 | 0.46395 | 1.3235

## Table 6
Table 5. Prevalence of STIs that were screened using multiplex real-time PCR TaqMan probe
No. | Patient's code | Code lab | Full name | Born | Age | Real-time PCR by "in house" kit (Ct) | [8 empty] | Leukocyte densities | Result
[6 empty] | G. vaginalis | N. gornohoeae | C. trachomatis | C. albicans | T. vaginalis | HSV-1 | HSV-2 | M. hominis | M. genitalium |  |  | No. | Abbreviation name of pathogen | n | % | Type | n | %
1 | 22027400 | HOG _01 | Luong T.T | 1995 | 27 | 32.99 |  |  | 28.74 | [5 empty] | ++ | GV+CA | 1 | GV | 181 | 33.83% | Single | 202 | 37.76%
2 | 16085090 | HOG _02 | Nguyen T.T.T | 1994 | 28 |  |  |  | 31.13 | [5 empty] | +++ | CA | 2 | NG | 1 | 0.1869% | Dual | 248 | 46.36%
3 | 21062583 | HOG _03 | Do T.H | 1998 | 24 | 36.79 |  |  | 37.00 | [5 empty] | ++ | GV+CA | 3 | CA | 19 | 3.551% | Triple | 63 | 11.78%
4 | 22027364 | HOG _11 | Ngo K.V | 2004 | 18 | 35.70 |  | 2022-02-24 | 34.51 | [5 empty] | +++ | GV+CT+CA | 4 | MH | 1 | 0.1869% | Quaq | 16 | 2.991%
5 | 22026378 | HOG _14 | Pham T.P | 1985 | 37 | 2022-05-21 |  | 36.25 | 32.0 |  |  |  | 35.85 |  | +++ | GV+CT+CA+MH | 5 | GV+CA | 204 | 38.13% | Pen | 3 | 0.5607%
6 | 17124031 | HOG _15 | Tran T.V.A | 1984 | 38 | 33.51 |  |  | 35.24 | [5 empty] | +++ | GV+CA | 6 | GV+NG | 1 | 0.1869% | Neg | 3 | 0.5607%
7 | 22026059 | HOG _26 | Dao T.L | 1994 | 28 | 23.04 | [8 empty] | +++ | GV | 7 | GV+CT | 20 | 3.738% | Total | 535 | 100%
8 | 22026060 | HOG _27 | Le T.T | 1979 | 43 | 29,37 | [8 empty] | +++ | GV | 8 | GV+TV | 2 | 0.3738% | CA | 295 | 55.14%
9 | 22026057 | HOG _28 | Nguyen T.T | 1983 | 39 | 37.43 | [8 empty] | +++ | GV | 9 | GV+HSV-1 | 2 | 0.3738% | GV | 503 | 94.02%
10 | 22026058 | HOG _29 | Mai T.T.H | 1981 | 41 |  |  |  | 24.70 | [5 empty] | ++ | CA | 10 | GV+MH | 11 | 2.056%
11 | 21025870 | HOG _30 | Nguyen T.M | 1997 | 25 | 23 |  |  | 27.59 | [5 empty] | ++ | GV+CA | 11 | NG+CA | 2 | 0.3738% | Abr name | n | %
12 | 22024798 | HOG _31 | Tran T.T | 1976 | 46 | 20.94 |  |  | 30.50 | [5 empty] | ++ | GV+CA | 12 | CA+MH | 4 | 0.7477% | GV | 503 | 94.02%
13 | 19004947 | HOG _32 | Nguyen T.H | 1995 | 27 | 17,37 | [8 empty] | +++ | GV | 13 | CA+MG | 2 | 0.3738% | NG | 25 | 4.673%
14 | 22024799 | HOG _33 | Nguyen T.T.H | 1999 | 23 | 21,59 | [8 empty] | +++ | GV | 14 | GV+NG+CT | 6 | 1.121% | CT | 66 | 12.34%
15 | 21024219 | HOG _34 | Nguyen T.Q.T | 1992 | 30 | 16.27 |  |  | 29.68 | [5 empty] | ++ | GV+CA | 15 | GV+CT+CA | 19 | 3.551% | CA | 295 | 55.14%
16 | 22018869 | HOG _35 | Pham T.T | 1999 | 23 | 20.22 |  |  | 28.28 | [5 empty] | ++ | GV+CA | 16 | GV+CT+HSV-1 | 1 | 0.1869% | TV | 7 | 1.308%
17 | 17007401 | HOG _37 | Nguyen T.T | 1989 | 33 | 25,53 | [8 empty] | +++ | GV | 17 | GV+CT+HSV-2 | 1 | 0.1869% | MH | 55 | 10.28%
18 | 22025579 | HOG _39 | Hoang T.K.P | 1997 | 25 | 25,64 | [8 empty] | nomal | GV | 18 | GV+NG+CA | 5 | 0.9346% | MG | 5 | 0.9346%
19 | 11030628 | HOG _41 | Pham T.N | 1984 | 38 |  |  |  | 29.14 | [5 empty] | ++ | CA | 19 | GV+CT+MH | 4 | 0.7477% | HSV-1 | 7 | 1.308%
20 | 20050103 | HOG _42 | Nguyen T.T | 1995 | 27 |  |  |  | 38.28 | [5 empty] | +++ | CA | 20 | GV+CA+HSV-1 | 2 | 0.3738% | HSV-2 | 3 | 0.5607%
21 | 20065451 | HOG _43 | Lam T.H | 2001 | 21 | 28.44 |  |  | 27.60 | [5 empty] | +++ | GV+CA | 21 | GV+CA+TV | 2 | 0.3738% | Total | 535 | 100%
22 | 14063813 | HOG_44 | Le T.H | 1994 | 28 | 17.72 |  |  | 33.58 |  |  |  | 21.82 |  | +++ | GV+CA+MH | 22 | GV+CA+MH | 21 | 3.925%
23 | 20056268 | HOG _45 | Vu T.H.T | 1992 | 30 | 18,23 | [8 empty] | +++ | GV | 23 | GV+CA+MG | 1 | 0.1869%
24 | 21079539 | HOG _46 | Dao N.H | 1996 | 26 | 16,35 | [8 empty] | +++ | GV | 24 | GV+TV+MH | 1 | 0.1869%
25 | 21058088 | HOG _47 | Do M.N | 1996 | 26 | 17.39 |  |  | 34.47 |  |  |  | 19.25 |  | +++ | GV+CA+MH | 25 | GV+NG+CT+CA | 4 | 0.7477%
26 | 20101744 | HOG _48 | Pham T.T | 1966 | 56 | 28.57 |  |  | 37.83 | [5 empty] | +++ | GV+CA | 26 | GV+NG+CT+HSV-1 | 1 | 0.1869%
27 | 17150225 | HOG _50 | Dao T.L | 1972 | 50 | 20.23 |  |  | 29.58 | [5 empty] | ++ | GV+CA | 27 | GV+NG+CT+MH | 1 | 0.1869%
28 | 22026029 | HOG _51 | Ho T.T | 1997 | 25 | 19.31 |  |  | 30.87 | [5 empty] | +++ | GV+CA | 28 | GV+NG+CA+MH | 1 | 0.1869%
29 | 16098791 | HOG _52 | Trinh T.X | 1991 | 31 |  |  |  | 37.70 |  |  |  |  | 29.62 | +++ | CA+MG | 29 | GV+CT+CA+MH | 6 | 1.121%
30 | 20072032 | HOG _53 | Tran D.V | 1992 | 30 | 16.65 |  |  | 33.43 | [5 empty] | +++ | GV+CA | 30 | GV+CA+HSV-1+MH | 1 | 0.1869%
31 | 20005321 | HOG _54 | Nguyen T.U | 1990 | 32 | 16.74 |  |  | 29.25 |  |  |  | 19.25 |  | +++ | GV+CA+MH | 31 | GV+TV+HSV-2+MH | 1 | 0.1869%
32 | 22024822 | HOG _55 | Pham T.Q | 1981 | 41 | 2022-11-18 |  |  | 31.27 |  |  |  | 18.81 |  | ++ | GV+CA+MH | 32 | GV+TV+MH+MG | 1 | 0.1869%
33 | 11046917 | HOG _56 | Nguyen T.T.N | 1987 | 35 | 34,71 | [8 empty] | +++ | GV | 33 | GV+NG+CT+CA+MH | 2 | 0.3738%
34 | 22027579 | HOG _57 | Hoang D.H | 2003 | 19 | 14.98 |  |  | 30.20 |  |  |  | 19.77 |  | +++ | GV+CA+MH | 34 | GV+NG+CT+HSV-2+MG | 1 | 0.1869%
35 | 22003026 | HOG _58 | Vu T.M.N | 1995 | 27 | 2022-08-29 |  |  | 36.68 |  |  |  | 28.54 |  | +++ | GV+CA+MH | 35 | Negative | 3 | 0.5607%
36 | 22027585 | HOG _59 | Tran T.P.L | 1980 | 42 | 34.79 |  |  | 29.23 | [5 empty] | +++ | GV+CA |  | Total | 535 | 100%
37 | 22027561 | HOG _61 | Nguyen T.T.H | 2003 | 19 | 34,11 | [8 empty] | +++ | GV
38 | 19154333 | HOG _62 | Do H.L | 1996 | 26 | 26,18 | [8 empty] | ++ | GV
39 | 22021599 | HOG _63 | Hoang T.T | 2000 | 22 | 31,13 | [8 empty] | ++ | GV
40 | 12085638 | HOG _64 | Vuong T.T | 1987 | 35 | 33.00 |  |  | 29.14 | [5 empty] | +++ | GV+CA
41 | 22024826 | HOG _65 | Nguyen T.T | 2000 | 22 | 2022-12-14 |  |  | 29.41 | [5 empty] | +++ | GV+CA
42 | 21064596 | HOG _67 | Dang T.B.N | 1999 | 23 | 20.88 |  |  | 27.24 | [5 empty] | +++ | GV+CA
43 | 22024435 | HOG _70 | Cao T.H.Y | 1993 | 29 | 13.67 |  |  | 28.29 | [5 empty] | +++ | GV+CA
44 | 22021597 | HOG _71 | Vu T.H.Y | 1989 | 33 | 19.69 |  |  | 27.84 | [5 empty] | +++ | GV+CA
45 | 12003709 | HOG _72 | Nguyen T.T.H | 1988 | 34 | 22.72 |  |  | 25.52 |  |  |  | 29.37 |  | ++ | GV+CA+MH
46 | 12087114 | HOG _73 | Dinh T.M.H | 1985 | 37 | 34.37 |  |  | 29.76 | [5 empty] | +++ | GV+CA
47 | 22016444 | HOG _75 | Nguyen T.T.H | 1998 | 24 | 13.94 |  |  | 26.81 | [5 empty] | +++ | GV+CA
48 | 22024869 | HOG _76 | Vu H.T | 1991 | 31 | 36.46 |  | 36.54 | 29.69 | [5 empty] | +++ | GV+CT+CA
49 | 18008374 | HOG _77 | Pham T.H | 1971 | 51 | 15.71 |  |  | 23.26 | [5 empty] | +++ | GV+CA
50 | 22054120 | HOG _78 | La T.T L | 1969 | 53 | 21,57 | [8 empty] | +++ | GV
51 | 22027835 | HOG _79 | Pham T.H | 1999 | 23 | 14,38 | [8 empty] | +++ | GV
52 | 15062758 | HOG _80 | Nguyen T.T.L | 1997 | 25 | 27,49 | [8 empty] | ++ | GV
53 | 22027775 | HOG _81 | Nguyen T.H | 1988 | 34 | 32.69 |  |  | 28.51 | [5 empty] | +++ | GV+CA
54 | 22018152 | HOG _82 | Vu T.H | 1990 | 32 | 30,26 | [8 empty] | +++ | GV
55 | 22023405 | HOG _83 | Tran T.L.D | 2000 | 22 | 16.16 |  |  | 2022-02-24 | [5 empty] | +++ | GV+CA
56 | 22027599 | HOG _84 | Nguyen T.D | 1977 | 45 | 27,15 | [8 empty] | +++ | GV
57 | 22024855 | HOG _86 | Ngo T.K.T | 1972 | 50 | 2022-02-21 | [8 empty] | ++ | GV
58 | 22027621 | HOG _87 | Ngo T.T | 1992 | 30 | 32,78 | [8 empty] | ++ | GV
59 | 22024863 | HOG _88 | Vu T.K | 1959 | 63 | 34,62 | [8 empty] | ++ | GV
60 | 22021594 | HOG _89 | Pham T.X.H | 1982 | 40 | 36,10 | [8 empty] | ++ | GV
61 | 22018703 | HOG _90 | Nguyen T.B | 1998 | 24 | 35.00 |  |  | 27.60 | [5 empty] | +++ | GV+CA
62 | 22027632 | HOG _92 | Nguyen T.L | 1982 | 40 | 2022-03-13 |  |  | 26.90 | [5 empty] | ++ | GV+CA
63 | 22024939 | HOG _93 | Nguyen T.M | 1963 | 59 | 15.84 |  |  | 32.59 | [5 empty] | +++ | GV+CA
64 | 22010308 | HOG _94 | Truong T.T.L | 1998 | 24 | 17,74 | [8 empty] | +++ | GV
65 | 22028374 | HOG _95 | Ngo H.Y | 1990 | 32 | 33,28 | [8 empty] | +++ | GV
66 | 22024945 | HOG _97 | Nguyen T.H.N | 1993 | 29 | 15.19 |  | 32.6 | 30.32 | [5 empty] | +++ | GV+CT+CA
67 | 22027992 | HOG _100 | Duong T.K | 1983 | 39 | 34,29 | [8 empty] | ++ | GV
68 | 22028385 | HOG _103 | Nguyen T.M.L | 1997 | 25 | 16,01 | [8 empty] | +++ | GV
69 | 22026558 | HOG _104 | Nguyen T.T | 2003 | 19 | 16.27 |  |  | 28.43 | [5 empty] | ++ | GV+CA
70 | 17080616 | HOG _105 | Dang T.H | 1988 | 34 | 14.6 |  |  | 28.07 | [5 empty] | +++ | GV+CA
71 | 22024944 | HOG _109 | Vu T.T | 1973 | 49 | 34.87 |  |  | 31.06 | [5 empty] | +++ | GV+CA
72 | 22024460 | HOG _110 | Hoang T.T.D | 1999 | 23 | 26,29 | [8 empty] | +++ | GV
73 | 21021193 | HOG _111 | Tran T.N | 1996 | 26 | 14.08 |  |  | 26..44 | [5 empty] | +++ | GV+CA
74 | 17030113 | HOG _112 | Trinh T.T.H | 1984 | 38 | 32.25 |  | 37.12 | 29.81 | [5 empty] | ++ | GV+CT+CA
75 | 15144115 | HOG _113 | Nguyen T.Q | 1988 | 34 | 13.55 |  |  | 32.09 | [5 empty] | +++ | GV+CA
76 | 16101606 | HOG _114 | Nguyen T.N | 1990 | 32 | 31,85 | [8 empty] | ++ | GV
77 | 17099362 | HOG _115 | Tran T.L | 1978 | 44 | 18,92 | [8 empty] | ++ | GV
78 | 15129238 | HOG _116 | Vu T.H | 2000 | 22 | 31.89 |  |  | 27.86 | [5 empty] | +++ | GV+CA
79 | 16153227 | HOG _117 | Tran T.P.T | 1991 | 31 | 22.33 |  |  | 28.41 | [5 empty] | +++ | GV+CA
80 | 22026553 | HOG _118 | Tran K.U | 2002 | 20 | 30.24 |  |  | 32.32 |  |  |  | 27.66 |  | ++ | GV+CA+MH
81 | 22022826 | HOG _119 | Nguyen T.H.Y | 2002 | 20 | 15.1 |  |  | 26.93 | [5 empty] | +++ | GV+CA
82 | 22026914 | HOG _120 | Do T.K.O | 2002 | 20 | 32.01 |  |  | 28.98 | [5 empty] | +++ | GV+CA
83 | 50201394 | HOG _121 | Ngo T.T.B | 1980 | 42 | 13.65 |  |  | 29.84 |  |  |  | 20.02 |  | +++ | GV+CA+MH
84 | 11003579 | HOG _122 | Hoang T.H | 1986 | 36 | 32,35 |  |  | - | [5 empty] | +++ | GV
85 | 22100220 | HOG _123 | Duong K.L | 2000 | 22 | 22.15 |  |  | 26.25 | [5 empty] | +++ | GV+CA
86 | 22027997 | HOG _124 | Pham T.D | 1990 | 32 | 20.55 |  |  | 30.72 | [5 empty] | +++ | GV+CA
87 | 22029399 | HOG _125 | Luc T.H | 1996 | 26 | 20,67 | [8 empty] | nomal | GV
88 | 22029501 | HOG _127 | Phan T.N | 1993 | 29 | 32,28 | [8 empty] | nomal | GV
89 | 12131976 | HOG _128 | Le T.U.R | 1983 | 39 | 31.73 | [8 empty] | nomal | GV
90 | 22028421 | HOG _129 | Nguyen T.N | 2002 | 20 | 33.61 |  | 36.16 | [6 empty] | nomal | GV+CT
91 | 22028426 | HOG _130 | Nong T.A.N | 2001 | 21 | 20.9 |  | 30.12 |  |  |  |  | 25.55 |  | nomal | GV+CT+MH
92 | 19064348 | HOG _131 | Quang T.H | 1995 | 27 | 25.43 | [8 empty] | +++ | GV
93 | 20045582 | HOG _132 | Pham T.Q.C | 1994 | 28 | 14.88 | [8 empty] | +++ | GV
94 | 22028702 | HOG _133 | Nguyen T.D | 1990 | 32 | 35.5 |  |  | 34.04 | [5 empty] | +++ | GV+CA
95 | 22025204 | HOG _134 | Lai N.Q | 2001 | 21 | 33.23 |  |  | 26.73 | [5 empty] | +++ | GV+CA
96 | 7057149 | HOG _135 | Vu B.H | 1982 | 40 | 36.31 |  |  | 27.56 | [5 empty] | +++ | GV+CA
97 | 8093275 | HOG _136 | Nguyen T.C | 1986 | 36 | 14.47 |  |  | 33.65 | [5 empty] | +++ | GV+CA
98 | 21006539 | HOG _137 | Tran T.T | 1989 | 33 | 35.11 | [8 empty] | ++ | GV
99 | 22015002 | HOG _138 | Doan T.H.T | 1995 | 27 | 17.45 |  |  | 29.62 | [5 empty] | +++ | GV+CA
100 | 14011445 | HOG _139 | Nguyen K.L | 1989 | 33 | 30.19 | [8 empty] | +++ | GV
101 | 14138977 | HOG _140 | Phung T.M.T | 1994 | 28 | 22.07 | [8 empty] | ++ | GV
102 | 7016833 | HOG _141 | Nguyen T.M | 1984 | 38 | 23.77 |  |  |  | 24.12 |  |  |  |  | +++ | GV+TV
103 | 19119781 | HOG _142 | Ta T.D | 1986 | 36 | 14.63 |  | 33.63 | 28.15 |  |  |  | 18.41 |  | +++ | GV+CT+CA+MH
104 | 22028464 | HOG _143 | Hoang T.A | 1998 | 24 | 14.78 |  |  | 27.63 | [5 empty] | +++ | GV+CA
105 | 22024091 | HOG _144 | Dao N.Q | 2001 | 21 | 31.41 |  | 19.76 | [6 empty] | ++ | GV+CT
106 | 22028683 | HOG _145 | Nguyen T.Q | 1999 | 23 | 23.83 |  |  | 35.64 | [5 empty] | +++ | GV+CA
107 | 22028685 | HOG _146 | Le H.Y | 2004 | 18 | 25.88 |  |  | 29.02 | [5 empty] | +++ | GV+CA
108 | 20078206 | HOG _147 | Nguyen T.N | 1996 | 26 | 15.57 |  | 35.36 | 29.21 | [5 empty] | +++ | GV+CT+CA
109 | 22026588 | HOG _148 | Pham T.H | 1987 | 35 | 27.91 |  |  | 28.81 | [5 empty] | ++ | GV+CA
110 | 22027923 | HOG _149 | Le L.A | 1992 | 30 | 33.83 |  |  | 29.18 | [5 empty] | +++ | GV+CA
111 | 12039912 | HOG _150 | Nguyen T.T | 1986 | 36 | 15.57 |  |  | 32.11 | [5 empty] | +++ | GV+CA
112 | 14075816 | HOG _151 | Do T.T | 1989 | 33 | 20.66 |  |  | 28.17 | [5 empty] | ++ | GV+CA
113 | 17050140 | HOG _153 | Dao T.M.A | 1993 | 29 | 14.14 | [8 empty] | +++ | GV
114 | 22021988 | HOG _155 | Nguyen D.L | 2000 | 22 | 14.77 |  | 21.58 | [6 empty] | ++ | GV+CT
115 | 22024760 | HOG _156 | Nguyen T.H | 2001 | 21 | 17.28 | [8 empty] | +++ | GV
116 | 21066715 | HOG _157 | Dinh T.H | 1975 | 47 | 15.41 |  |  | 31.77 | [5 empty] | ++ | GV+CA
117 | 19047029 | HOG _158 | Le T.T.T | 1995 | 27 | 28.71 | [8 empty] | ++ | GV
118 | 15110465 | HOG _159 | Nguyen H.L | 1991 | 31 | 14.85 |  |  | 28.64 | [5 empty] | ++ | GV+CA
119 | 22024488 | HOG _160 | Pham T.T.H | 2003 | 19 | 32.48 |  |  | 32.25 | [5 empty] | ++ | GV+CA
120 | 22028942 | HOG _161 | Nguyen T.T.T | 1989 | 33 | 15.78 | [8 empty] | + | GV
121 | 22028935 | HOG _162 | Nguyen T.P.T | 2001 | 21 | 33.67 |  |  | 31.33 | [5 empty] | +++ | GV+CA
122 | 22028948 | HOG _163 | Nguyen T.T.L | 1974 | 48 | 32.29 |  |  | 28.72 | [5 empty] | +++ | GV+CA
123 | 22028952 | HOG _164 | Truong T.H | 1991 | 31 | 16.23 |  |  | 30.41 |  |  |  | 16.63 |  | ++ | GV+CA+MH
124 | 12085798 | HOG _165 | Nguyen T.B..D | 1983 | 39 | 15.76 |  |  | 26.27 | [5 empty] | +++ | GV+CA
125 | 22028527 | HOG _166 | Dang T.M.N | 1988 | 34 | 31.54 | [8 empty] | ++ | GV
126 | 22028499 | HOG _167 | To T.R | 1993 | 29 | 16.31 |  |  | 37.27 | [5 empty] | +++ | GV+CA
127 | 19632574 | HOG _168 | Le D.V | 1990 | 32 | 29.3 | [8 empty] | ++ | GV
128 | 22028502 | HOG _169 | Tran T.K.A | 1993 | 29 | 30.05 |  |  | 27.59 | [5 empty] | ++ | GV+CA
129 | 18270196 | HOG _170 | Vu T.D | 1998 | 24 | 31.84 | [8 empty] | ++ | GV
130 | 11194980 | HOG _171 | Do T.T | 1990 | 32 | 19.24 |  |  | 28.56 | [5 empty] | ++ | GV+CA
131 | 22029024 | HOG _172 | Nguyen T.H | 1993 | 29 | 15.57 |  |  | 31.41 | [5 empty] | ++ | GV+CA
132 | 21069689 | HOG _173 | Nguyen T.T | 2000 | 22 | 33.24 | [8 empty] | nomal | GV
133 | 22029541 | HOG _174 | Ta T.T | 1991 | 31 | 31.58 | [8 empty] | nomal | GV
134 | 19140093 | HOG _175 | Nguyen T.L | 1994 | 28 | 19.77 |  | 28.3 | [7 empty] | GV+CT
135 | 22028574 | HOG _180 | Pham T.P.T | 1998 | 24 | 17.93 |  |  | 27.99 | [5 empty] | +++ | GV+CA
136 | 22028881 | HOG _181 | Nguyen T.Q | 1996 | 26 |  |  |  | 27.68 | [5 empty] | +++ | CA
137 | 17104610 | HOG _182 | Nguyen T.H | 1993 | 29 | 14.62 |  |  | 27.51 | [5 empty] | +++ | GV+CA
138 | 21048600 | HOG _183 | Nguyen H.H | 1992 | 30 | 17.09 |  |  | 27.32 | [5 empty] | +++ | GV+CA
139 | 22028578 | HOG _184 | Vu T.T.T | 1988 | 34 | 16.58 |  |  | 28.08 | [5 empty] | +++ | GV+CA
140 | 11063728 | HOG _185 | Ha N.T | 1992 | 30 | 23.79 | 37.75 |  | 30.13 |  |  |  | 33.11 |  | +++ | GV+NG+CA+MH
141 | 22026229 | HOG _187 | Nguyen T.M | 1995 | 27 | 30.82 | [8 empty] | +++ | GV
142 | 22028871 | HOG _188 | Mai T.C | 2003 | 19 |  |  |  | 27.14 | [5 empty] | +++ | CA
143 | 22026639 | HOG _189 | Nguyen T.T.T.H | 1973 | 49 | 35.33 | [8 empty] | ++ | GV
144 | 13102188 | HOG _190 | Hoang T.T.L | 1989 | 33 | 27.39 |  | 27.54 | [6 empty] | ++ | GV+CT
145 | 22028606 | HOG _191 | Nguyen T.T | 1989 | 33 | 17.21 |  |  | 29.05 | [5 empty] | +++ | GV+CA
146 | 22027695 | HOG _192 | Do T.P | 1992 | 30 | 21.18 | [8 empty] | +++ | GV
147 | 12143074 | HOG _194 | Do T.L | 1991 | 31 | 27.56 | [8 empty] | ++ | GV
148 | 19135315 | HOG _195 | Vu T.D | 1999 | 23 | 21.69 | [8 empty] | ++ | GV
149 | 19007296 | HOG _196 | Ma T.H | 1994 | 28 | 35.67 | [8 empty] | +++ | GV
150 | 21056378 | HOG _198 | Nguyen T.H | 1992 | 30 | 22.59 | [8 empty] | ++ | GV
151 | 22028858 | HOG _199 | Nguyen T.T.L | 1999 | 23 | 34.7 | [8 empty] | +++ | GV
152 | 22028549 | HOG _200 | Hoang T.D | 1982 | 40 | 14.61 |  |  | 29.64 | [5 empty] | +++ | GV+CA
153 | 22029636 | HOG _202 | Dao T.H | 1999 | 23 | 19.3 | [8 empty] | nomal | GV
154 | 12139840 | HOG _203 | Phan T.H | 1989 | 33 | 34.57 | [8 empty] | nomal | GV
155 | 22030343 | HOG _204 | Vu T.T | 1972 | 50 | 35.57 | [6 empty] | 27.42 |  | +++ | GV+MH
156 | 22027900 | HOG _205 | Vu T.L.A | 1986 | 36 | 34.73 | [8 empty] | +++ | GV
157 | 22027891 | HOG _206 | Tran T.T.L | 2000 | 22 | 34.69 |  |  | 32.5 | [5 empty] | +++ | GV+CA
158 | 22028640 | HOG _207 | Nguyen N.H | 1990 | 32 | 16.59 |  |  | 28.42 | [5 empty] | +++ | GV+CA
159 | 22029057 | HOG _208 | Nguyen T.K.H | 1987 | 35 | 16.32 |  |  | 30.47 | [5 empty] | +++ | GV+CA
160 | 22026664 | HOG _209 | Doan T.H | 1995 | 27 | 33.97 |  |  | 30.78 | [5 empty] | +++ | GV+CA
161 | 21018278 | HOG _210 | Nguyen T.P | 1983 | 39 | 18.02 |  |  | 27.98 | [5 empty] | +++ | GV+CA
162 | 15127563 | HOG _211 | Pham T.T | 1988 | 34 | 36.06 | [8 empty] | +++ | GV
163 | 16017019 | HOG _212 | Nguyen P.H | 1990 | 32 | 35.05 | [8 empty] | ++ | GV
164 | 14023794 | HOG _213 | Nguyen T.H | 1970 | 52 | 18.36 | [8 empty] | ++ | GV
165 | 22024508 | HOG _214 | Hoang P.U | 2001 | 21 | 35.17 | [8 empty] | +++ | GV
166 | 17110743 | HOG _215 | Mai T.T.N | 1993 | 29 | 16.07 |  |  | 29.41 | [5 empty] | +++ | GV+CA
167 | 51000647 | HOG _216 | Nguyen T.N | 1980 | 42 | 32.49 |  | 39.32 | [6 empty] | ++ | GV
168 | 22010672 | HOG _217 | Nguyen T.H.G | 2001 | 21 | 29.44 | 36.12 | 36.94 | [6 empty] | +++ | GV+NG+CT
169 | 22030341 | HOG _218 | Vu M.T | 1995 | 27 | 16.94 |  |  | 29.32 | [5 empty] | ++ | GV+CA
170 | 22031368 | HOG _219 | Nguyen T.T | 1995 | 27 | 25.81 | [8 empty] | +++ | GV
171 | 22030338 | HOG _220 | Thach K.A | 1977 | 45 | 26.36 |  | 38.24 |  |  | 36.8 |  |  |  | ++ | GV+CT+HSV-1
172 | 22029620 | HOG _222 | Nguyen T.T.T.H | 1991 | 31 | 15.71 |  |  | 30.62 | [5 empty] | +++ | GV+CA
173 | 41005220 | HOG _223 | Nguyen T.N | 1985 | 37 | 33.18 | [8 empty] | +++ | GV
174 | 22028996 | HOG _224 | Pham T.N | 2002 | 20 | 14.72 | [6 empty] | 26.92 |  | +++ | GV+MH
175 | 19129382 | HOG _226 | Nguyen T.N | 1996 | 26 | 35.85 | [8 empty] | +++ | GV
176 | 14074892 | HOG _227 | Nguyen T.H | 1989 | 33 | 27.93 | [8 empty] | +++ | GV
177 | 17007961 | HOG _228 | Trinh T.T.T | 1990 | 32 |  |  |  | 27.88 | [5 empty] | ++ | CA
178 | 18042367 | HOG _229 | Nguyen T.T.N | 1991 | 31 |  |  |  | 29.2 | [5 empty] | +++ | CA
179 | 12007018 | HOG _231 | Do T.C | 1986 | 36 |  |  |  | 26.34 | [5 empty] | +++ | CA
180 | 20024053 | HOG _232 | Nguyen T.D.H | 1985 | 37 |  |  |  | 28.24 | [5 empty] | +++ | CA
181 | 21024555 | HOG _233 | Nguyen T.T.L | 1985 | 37 | 17.26 | [8 empty] | +++ | GV
182 | 16035505 | HOG _234 | Ta T.H | 1989 | 33 | 17.68 |  |  |  |  | 12.8 |  |  |  | +++ | GV+HSV-1
183 | 22030334 | HOG _235 | Pham T.H | 1977 | 45 | 29.53 | 37.04 | 38.01 |  |  | 36.49 |  |  |  | +++ | GV+NG+CT+HSV-1
184 | 22028661 | HOG _236 | Nguyen T.K.H | 1992 | 30 | 22.4 |  |  | 34.33 | [5 empty] | ++ | GV+CA
185 | 40707187 | HOG _237 | Dinh T.P | 1982 | 40 | 37.6 |  |  | 29.35 | [5 empty] | +++ | GV+CA
186 | 22031382 | HOG _239 | Nguyen T.T | 1977 | 45 | 21.24 | [8 empty] | ++ | GV
187 | 22030729 | HOG _240 | Nguyen T.T | 1995 | 27 | 33.19 | [8 empty] | ++ | GV
188 | 21017313 | HOG _241 | Thai T.A | 1991 | 31 | 27.77 |  |  |  |  | 37.67 |  |  |  | +++ | GV+HSV-1
189 | 22024190 | HOG _242 | Hoang T.T | 1976 | 46 | 28.91 |  |  | 31.22 |  | 36.66 |  |  |  | +++ | GV+CA+HSV-1
190 | 22026014 | HOG _243 | Vu T.T | 2001 | 21 | 22.92 | [8 empty] | +++ | GV
191 | 22030731 | HOG _244 | Nguyen T.K.T | 2002 | 20 | 31.82 |  |  | 31.11 | [5 empty] | ++ | GV+CA
192 | 16023572 | HOG _245 | Vu T.H | 1991 | 31 | 17.12 |  |  | 29.45 |  | 35.36 |  |  |  | +++ | GV+CA+HSV-1
193 | 22028045 | HOG _246 | Mai T.H | 1996 | 26 | 29.1 | [8 empty] | ++ | GV
194 | 13027564 | HOG _247 | Trinh T.T.H | 1987 | 35 | 35.94 |  |  | 30.14 | [5 empty] | +++ | GV+CA
195 | 22030779 | HOG _248 | Tran T.T | 2000 | 22 | 14.78 | [8 empty] | ++ | GV
196 | 21020511 | HOG _249 | Cao T.K.L | 1988 | 34 | 27.7 |  |  | 30.56 | [5 empty] | ++ | GV+CA
197 | 22030780 | HOG _250 | Duong H.A | 2003 | 19 |  |  |  | 27.9 | [5 empty] | +++ | GV+CA
198 | 16152319 | HOG _251 | Pham B.H | 1987 | 35 | 20.25 |  |  | 29.07 | [5 empty] | +++ | GV+CA
199 | 22004147 | HOG _252 | Pham T.N | 1988 | 34 | 14.35 |  |  | 30.03 | [5 empty] | +++ | GV+CA
200 | 16138779 | HOG _253 | Nguyen K.H | 1997 | 25 | 33.08 |  |  | 27.75 |  |  |  | 38.23 |  | +++ | GV+CA+MH
201 | 7046864 | HOG _254 | Le T.X | 1981 | 41 | 16.46 | [8 empty] | +++ | GV
202 | 22031508 | HOG _255 | Xa T.T | 2001 | 21 | 14.82 |  | 31.07 | 29 | [5 empty] | +++ | GV+CT+CA
203 | 22028220 | HOG _256 | Do T.M | 1978 | 44 | 15.24 | [8 empty] | ++ | GV
204 | / | HOG _257 | Nguyen T.G | 1997 | 25 | 33.91 |  |  | 28.82 | [5 empty] | +++ | GV+CA
205 | 10115801 | HOG _258 | Chu T.H | 1984 | 38 | 14.35 |  |  | 25.84 | [5 empty] | +++ | GV+CA
206 | 22051234 | HOG _259 | Pham T.T.H | 1995 | 27 | 13.68 | [8 empty] | +++ | GV
207 | 11032204 | HOG _260 | Nguyen T.T | 1991 | 31 | 12.58 | [6 empty] | 22.11 |  | ++ | GV+MH
208 | 22020398 | HOG _261 | Nguyen T.H.Y | 1992 | 30 | 17.86 |  |  | 27.62 | [5 empty] | ++ | GV+CA
209 | 19126609 | HOG _262 | Ngo T.T | 1985 | 37 | 14.88 | [8 empty] | +++ | GV
210 | 22030796 | HOG _263 | Nguyen B.N | 1998 | 24 | 31 | [8 empty] | +++ | GV
211 | 20066026 | HOG _268 | Nguyen V.V.P | 1992 | 30 | 14.54 |  |  | 25.83 | [5 empty] | +++ | GV+CA
212 | 22030893 | HOG _269 | Nham T.T | 1995 | 27 | 17.37 |  |  | 24.48 | [5 empty] | ++ | GV+CA
213 | 22024397 | HOG _270 | Vu T.H | 1985 | 37 | 14.91 | [8 empty] | +++ | GV
214 | 22030874 | HOG _271 | Le T.K.O | 2000 | 22 | 13.84 | [8 empty] | +++ | GV
215 | 20048518 | HOG _272 | Ngo H.P | 1993 | 29 | 27.19 | [8 empty] | ++ | GV
216 | 20011337 | HOG _273 | Cao T.T | 1991 | 31 | 34.95 | [8 empty] | ++ | GV
217 | 7065450 | HOG _274 | Tran T.K.H | 1975 | 47 | 23.97 |  |  | 26.82 | [5 empty] | +++ | GV+CA
218 | 22030865 | HOG _275 | Le T.H.N | 2000 | 22 | 18.78 |  |  | 27.81 | [5 empty] | +++ | GV+CA
219 | 11030291 | HOG _276 | Tran P.A | 1979 | 43 | 25.37 | [8 empty] | ++ | GV
220 | 22027942 | HOG _277 | Dinh D.C | 2001 | 21 | 17.46 |  |  | 32.08 | [5 empty] | +++ | GV+CA
221 | 22030866 | HOG _278 | Vo T.N | 1990 | 32 | 23.89 |  |  | 24.99 | [5 empty] | +++ | GV+CA
222 | 22027941 | HOG _279 | Le T.L | 2000 | 22 | 35.56 | [8 empty] | ++ | GV
223 | 22030825 | HOG _280 | Nguyen T.P.H | 2001 | 21 | 35.42 | [8 empty] | +++ | GV
224 | 14050470 | HOG _281 | Phan T.H | 1985 | 37 | 33.53 | [8 empty] | +++ | GV
225 | 16100227 | HOG _282 | Can T.T.L | 1998 | 24 | 19.91 | [8 empty] | ++ | GV
226 | 51005000 | HOG _283 | Nguyen T.T.T | 1984 | 38 | 30.61 | [8 empty] | ++ | GV
227 | 15053137 | HOG _285 | Cao T.T | 1974 | 48 | 32.17 | [8 empty] | ++ | GV
228 | 12036867 | HOG _286 | Pham T.H,N | 1982 | 40 | 26.73 |  |  | 29.61 | [5 empty] | +++ | GV+CA
229 | 22030870 | HOG _287 | Trương T.H | 1970 | 52 | 32.09 |  |  | 26.01 | [5 empty] | +++ | GV+CA
230 | 18125742 | HOG _288 | Nguyen T.H | 1988 | 34 | 33.11 | [8 empty] | ++ | GV
231 | 22004550 | HOG _289 | Le T.N | 2002 | 20 | 16.39 |  |  | 27.68 | [5 empty] | ++ | GV+CA
232 | 22030837 | HOG _290 | Hoang T.H | 2000 | 22 |  |  |  | 27.72 |  |  |  | 21.81 |  | ++ | CA+MH
233 | 22010459 | HOG _291 | Nguyen T.P | 2002 | 20 |  |  |  | 30.34 |  |  |  | 37.64 |  | ++ | CA+MH
234 | 12081133 | HOG _292 | Nguyen T.G | 1990 | 32 | 15.68 | [6 empty] | 11.8 |  | ++ | GV+MH
235 | 22031011 | HOG _294 | Dinh T.H | 1999 | 23 | 19.25 |  |  | 31.57 | [5 empty] | nomal | GV+CA
236 | 22032004 | HOG _296 | Vu T.H | 1991 | 31 | 25.88 | [8 empty] | nomal | GV
237 | 22017128 | HOG _297 | Ngo H.A | 1998 | 24 | 14.25 |  |  | 27.98 | [5 empty] | +++ | GV+CA
238 | 20052184 | HOG _298 | Le T.T | 1988 | 34 | 20.94 |  | 23.73 | [6 empty] | +++ | GV+CA
239 | 13102188 | HOG _299 | Hoang T.T.L | 1989 | 33 | 31.52 |  | 24.49 | [6 empty] | +++ | GV+CT
240 | 22029808 | HOG _300 | Bui T.T.H | 1996 | 26 |  |  | 24.85 | 30.76 | [5 empty] | +++ | GV+CT+CA
241 | 20015420 | HOG _301 | Ngo T.B.N | 1993 | 29 | [9 empty] | +++ | Negative
242 | 22032000 | HOG _302 | Bui T.M | 1998 | 24 | 17.75 |  |  | 30.07 | [5 empty] | +++ | GV+CA
243 | 22030949 | HOG _303 | Nguyen T.T.A | 1989 | 33 | 14.32 |  |  | 25.26 | [5 empty] | +++ | GV+CA
244 | 17254711 | HOG _304 | Quach T.H | 1989 | 33 | 17.89 |  |  | 26.63 | [5 empty] | +++ | GV+CA
245 | 21073580 | HOG _305 | Kieu T.T.T | 1993 | 29 | 27.74 |  |  | 26.13 | [5 empty] | +++ | GV+CA
246 | 21067119 | HOG _306 | Do T.N | 1993 | 29 | 25.53 |  |  | 28 | [5 empty] | +++ | GV+CA
247 | 19067656 | HOG _307 | Ho T.H | 1993 | 29 | 28.5 |  |  | 28.03 | [5 empty] | +++ | GV+CA
248 | 18027184 | HOG _308 | Nguyen M.H | 1993 | 29 | 20.4 |  |  | 25.84 |  |  |  | 35.97 |  | +++ | GV+CA+MH
249 | 22030987 | HOG _309 | Ngo P.A | 2002 | 20 | 15.3 |  |  | 27.24 | [5 empty] | +++ | GV+CA
250 | 22012620 | HOG _310 | Nguyen Q.N | 1994 | 28 |  |  |  | 27.39 |  |  |  |  | 32.85 | +++ | CA+MG
251 | 22028193 | HOG _311 | Pham T. | 1988 | 34 |  |  |  | 28.06 | [5 empty] | +++ | CA
252 | 22030988 | HOG _312 | Do T.T.V | 1984 | 38 | 19.74 |  |  | 27.77 |  |  |  | 17.44 |  | +++ | GV+CA+MH
253 | 12053535 | HOG _314 | Ngo H.H | 1977 | 45 |  |  |  | 28.1 | [5 empty] | +++ | CA
254 | 22013847 | HOG _315 | Tran L.N | 1997 | 25 | 20.45 |  |  | 28.33 | [5 empty] | +++ | GV+CA
255 | 18106149 | HOG _316 | Nguyen T.L | 1994 | 28 | 13.28 |  |  | 28.11 | [5 empty] | +++ | GV+CA
256 | 21070105 | HOG _317 | Nguyen N.A | 2000 | 22 | 17.65 |  |  | 30.77 | [5 empty] | ++ | GV+CA
257 | 22030490 | HOG _318 | Le T.V.A | 2001 | 21 | 16.56 | [6 empty] | 25.43 |  | ++ | GV+MH
258 | 22032120 | HOG _319 | Nguyen T.T.T | 2003 | 19 | 20.29 |  |  | 28.11 | [5 empty] | +++ | GV+CA
259 | 22027708 | HOG _320 | Nguyen T.H | 2001 | 21 | 30.82 |  | 24.39 | 27.47 | [5 empty] | ++ | GV+CT+CA
260 | 22011071 | HOG _321 | Pham T.T.D | 1999 | 23 | 15.43 |  |  | 29.88 |  |  |  | 26.73 |  | ++ | GV+CA+MH
261 | 22032209 | HOG _323 | Nguyen T.H.Y | 1997 | 25 | 26.91 | [8 empty] | +++ | GV
262 | 22032208 | HOG _324 | Vuong T.Q.N | 1999 | 23 | 18.85 |  |  | 37.35 | [5 empty] | ++ | GV+CA
263 | 17015545 | HOG _325 | Nguyen T.H | 1995 | 27 | 32.91 |  |  | 25.98 | [5 empty] | ++ | GV+CA
264 | 21066288 | HOG _327 | Phan T.K.L | 2003 | 19 | 33.69 | [8 empty] | ++ | GV
265 | 19025142 | HOG _328 | Chu M.,N | 1997 | 25 | 36.65 |  |  | 38.74 | [5 empty] | ++ | GV+CA
266 | 21008570 | HOG _329 | Nguyen T.T | 1994 | 28 | 21.99 |  |  | 27.74 | [5 empty] | +++ | GV+CA
267 | 20260304 | HOG _330 | Duong T.B.N | 1989 | 33 | 30.07 |  |  | 29.81 | [5 empty] | ++ | GV+CA
268 | 22032137 | HOG _331 | Bui T.X | 1973 | 49 | 19.61 | [8 empty] | ++ | GV
269 | 22029721 | HOG _332 | Nuyen T.T | 1989 | 33 | 16.4 |  |  | 29.66 | [5 empty] | +++ | GV+CA
270 | 22030513 | HOG _335 | Do T.L.H | 1989 | 33 | 19.37 | [8 empty] | +++ | GV
271 | 14063103 | HOG _336 | Nguyen T.H | 1982 | 40 | 21.23 | [8 empty] | +++ | GV
272 | 16016857 | HOG _337 | Bui T.T.H | 1992 | 30 | 16.72 | [6 empty] | 24.44 |  | +++ | GV+MH
273 | 22031210 | HOG _338 | Kieu N.N | 1997 | 25 | 20.99 |  | 25.79 | [6 empty] | nomal | GV+CT
274 | 22031209 | HOG _339 | Vuương T.T.T | 1996 | 26 | 25.97 |  | 30.58 |  |  |  | 31.73 |  |  | nomal | GV+CT+HSV-2
275 | 7057739 | HOG _340 | Le V.H | 1981 | 41 | 31.83 | [8 empty] | nomal | GV
276 | 18274109 | HOG _341 | Nguyen T.T.N | 1992 | 30 | 26.05 | [8 empty] | nomal | GV
277 | 20100249 | HOG _342 | Vu T.V | 1995 | 27 | 24.55 | [8 empty] | nomal | GV
278 | 22031215 | HOG _343 | Nguyen V.T | 1996 | 26 | 23.83 |  | 25.87 | [6 empty] | nomal | GV+CT
279 | 22031158 | HOG _344 | Vu T.H | 1985 | 37 | 34.66 |  | 31.6 | [6 empty] | nomal | GV+CT
280 | 22031177 | HOG _346 | Ha H.T | 1992 | 30 | 20.6 | [8 empty] | nomal | GV
281 | 22031208 | HOG _347 | Nguyen T.O | 1989 | 33 | 30.04 | [8 empty] | nomal | GV
282 | 22031207 | HOG _348 | Nguyen T.H | 1999 | 23 | 32.23 | [8 empty] | nomal | GV
283 | 15024524 | HOG _352 | Nguyen T.H | 2002 | 20 | 35.49 | [8 empty] | +++ | GV
284 | 20068003 | HOG _355 | Nguyen T.L | 1996 | 26 | 17.22 | [8 empty] | +++ | GV
285 | 12031289 | HOG _357 | Lưu T.H | 1991 | 31 | 21.66 |  |  | 26.67 | [5 empty] | +++ | GV+CA
286 | 17080068 | HOG _358 | Nguyen P.L.T | 1996 | 26 | 18.01 |  |  | 29.1 | [5 empty] | +++ | GV+CA
287 | 22023459 | HOG _359 | Truong T.M | 2001 | 21 |  |  |  | 28.62 | [5 empty] | +++ | CA
288 | 19039788 | HOG _360 | Nguyen T.T.N | 1995 | 27 | 16.66 | [8 empty] | +++ | GV
289 | 19104280 | HOG _362 | Nguyen T.T.T | 1995 | 27 | 34.42 |  |  | 31.11 | [5 empty] | +++ | GV+CA
290 | 19129181 | HOG _363 | Hoang T.C | 1994 | 28 | 16.25 | [8 empty] | +++ | GV
291 | 19143857 | HOG _364 | Ngo T.T.H | 1993 | 29 | 17.77 |  |  | 28.45 | [5 empty] | +++ | GV+CA
292 | 22004565 | HOG _367 | Ngac T.T.L | 2001 | 21 | 36.23 | [8 empty] | ++ | GV
293 | 22034020 | HOG _368 | Do H.L | 1999 | 23 | 16.22 |  |  | 28.54 | [5 empty] | ++ | GV+CA
294 | 17039958 | HOG _369 | Ngo T.H.H | 1988 | 34 | 31.2 |  |  | 27.86 | [5 empty] | +++ | GV+CA
295 | 19046159 | HOG _370 | Tran T.C | 1991 | 31 | 15.46 | [8 empty] | +++ | GV
296 | 20005279 | HOG _371 | Doan T.N | 1978 | 44 | 36.37 |  |  | 28.45 | [5 empty] | ++ | GV+CA
297 | 22034021 | HOG _372 | Duong K.C | 1985 | 37 | 34.51 |  |  | 26.92 | [5 empty] | +++ | GV+CA
298 | 21053371 | HOG _373 | Le T.P | 1993 | 29 | 16.61 | [8 empty] | +++ | GV
299 | 22034014 | HOG _374 | Bui T.T.H | 1996 | 26 | 17.78 | [8 empty] | +++ | GV
300 | 20045530 | HOG _375 | Tran T.H | 1994 | 28 | 18.29 |  |  | 27.76 | [5 empty] | +++ | GV+CA
301 | 18053417 | HOG _376 | Nguyen T.T.H | 1976 | 46 | 35.4 | [8 empty] | +++ | GV
302 | 22031635 | HOG _377 | Nguyen T.Y | 1966 | 56 | 34.89 | [8 empty] | +++ | GV
303 | 22031630 | HOG _378 | Truong T.H | 1999 | 23 | 18.35 |  |  | 33.89 | [5 empty] | +++ | GV+CA
304 | 22034013 | HOG _379 | Le P.A | 2003 | 19 | 35.65 |  |  | 33.39 | [5 empty] | +++ | GV+CA
305 | 17020207 | HOG _380 | Do T.N | 1990 | 32 | 36.33 | [8 empty] | +++ | GV
306 | 13138891 | HOG _381 | Nguyen T.D | 1984 | 38 | 27.74 |  |  | 34.11 | [5 empty] | +++ | GV+CA
307 | 22032315 | HOG _382 | Vu T.L.A | 2000 | 22 | 16.22 |  |  | 28.43 | [5 empty] | ++ | GV+CA
308 | 22034000 | HOG _383 | Nguyen T.H | 2001 | 21 | 19.97 |  |  | 33.02 | [5 empty] | ++ | GV+CA
309 | 22033549 | HOG _384 | Dang T.T.C | 1982 | 40 | 18.84 |  | 21.76 |  |  |  |  | 36.79 |  | +++ | GV+CT+MH
310 | 9037769 | HOG _385 | Le K.O | 1985 | 37 | 22.69 | [8 empty] | +++ | GV
311 | 22031633 | HOG _386 | Ngo T.T.L | 1964 | 58 | 16.54 | [8 empty] | +++ | GV
312 | 18081249 | HOG _387 | Lam T.H | 1988 | 34 | 27.51 | [8 empty] | ++ | GV
313 | 22031622 | HOG _388 | Ngo T.T | 1966 | 56 | 29.37 | [8 empty] | ++ | GV
314 | 7096646 | HOG _389 | Nguyen T.P.D | 1982 | 40 | 28.39 | [8 empty] | ++ | GV
315 | 15093075 | HOG _391 | Luu T.T | 1987 | 35 | 18.87 | [8 empty] | +++ | GV
316 | 22033554 | HOG _392 | Duong T.H.N | 1997 | 25 | 17.16 | [8 empty] | +++ | GV
317 | 22032318 | HOG _393 | Dang T.H | 1978 | 44 | 31.38 | [8 empty] | ++ | GV
318 | 20079186 | HOG _394 | Cu T.K.H | 1997 | 25 | 19.15 |  |  | 28.71 | [5 empty] | +++ | GV+CA
319 | 22001486 | HOG _395 | Nguyen T.L | 1996 | 26 | 30.03 |  |  | 37.92 | [5 empty] | ++ | GV+CA
320 | 13126498 | HOG _396 | Tran K.L | 1990 | 32 | 16.19 |  |  | 27.48 | [5 empty] | +++ | GV+CA
321 | 22032299 | HOG _397 | Bui T.X' | 1968 | 54 | 22.63 |  |  | 36.68? |  |  |  | 25.93 |  | +++ | GV+MH
322 | 17130923 | HOG _398 | Le T.H.T | 1992 | 30 | 31.25 |  |  | 28.53 | [5 empty] | ++ | GV+CA
323 | 20057522 | HOG _399 | Nguyen K.L | 2001 | 21 | 33.69 |  |  | 32.43 |  |  |  |  | 31.17 | +++ | GV+CA+MG
324 | 21057645 | HOG _400 | Tran T.D | 1996 | 26 | 18.03 | [8 empty] | ++ | GV
325 | 22031609 | HOG _401 | Nguyen H.M | 2003 | 19 | 34.76 |  | 24.21 | [6 empty] | +++ | GV+CT
326 | 12060384 | HOG _402 | Nguyen T.M.T | 1977 | 45 | 32.16 |  |  | 26.9 | [5 empty] | +++ | GV+CA
327 | 16126952 | HOG _404 | Vu T.T | 1993 | 29 | 23.12 | [8 empty] | +++ | GV
328 | 18031536 | HOG _405 | Tran T.N | 1996 | 26 | 31.19 |  |  | 28.4 | [5 empty] | ++ | GV+CA
329 | 22032278 | HOG _406 | Le T.C | 1999 | 23 | 17.37 | [8 empty] | +++ | GV
330 | 19095825 | HOG _407 | Do T.N | 1998 | 24 | 32.34 |  |  | 27.78 | [5 empty] | ++ | GV+CA
331 | 8016547 | HOG _408 | Nguyen T.H | 1978 | 44 | 29.63 |  |  | 34.41 | [5 empty] | +++ | GV+CA
332 | 18118408 | HOG _409 | Nguyen T.T | 1995 | 27 | 32.87 |  |  | 30.91 |  |  |  | 24.67 |  | +++ | GV+CA+MH
333 | 22030613 | HOG _410 | Vu T.H.H | 1999 | 23 | 15.96 |  | 35.95 | 27.14 |  |  |  | 37.36 |  | +++ | GV+CT+CA+MH
334 | 22034032 | HOG _411 | Nguyen M.D.L | 2003 | 19 | 19.45 |  |  | 28.34 | [5 empty] | +++ | GV+CA
335 | 22024575 | HOG _412 | Chu T.M | 1994 | 28 | 20.85 |  | 35.83 | 28.76 | [5 empty] | +++ | GV+CT+CA
336 | 30802051 | HOG _413 | Trinh T.N.T | 1978 | 44 | 18.11 |  | 35 |  |  |  |  | 31.15 |  | +++ | GV+CT+MH
337 | 22033506 | HOG _415 | Nguyen T.T.H | 2002 | 20 | 33.9 |  |  | 28.49 | [5 empty] | +++ | GV+CA
338 | 22034085 | HOG _417 | Le T.H.L | 1972 | 50 | 35.38 | [8 empty] | +++ | GV
339 | 22030615 | HOG _418 | Nguyen T.H' | 2003 | 19 | 16.84 |  |  | 28.17 | [5 empty] | +++ | GV+CA
340 | 18061928 | HOG _420 | Le T.H | 1997 | 25 |  |  |  | 36.49 |  |  |  | 19.8 |  | ++ | CA+MH
341 | 13055180 | HOG _421 | Nguyen T.T | 1981 | 41 | 33.32 | [8 empty] | +++ | GV
342 | 22034036 | HOG _423 | Le T.T | 1987 | 35 | 34.39 |  |  | 30.33 | [5 empty] | ++ | GV+CA
343 | 13058799 | HOG _424 | Do T.H | 1991 | 31 | 18.97 |  |  | 32.9 | [5 empty] | +++ | GV+CA
344 | 8043938 | HOG _425 | Nguyen T.M.Y | 1974 | 48 | 35.32 | [8 empty] | ++ | GV
345 | 11170036 | HOG _426 | Tran T.P.T | 1986 | 36 | 35.41 |  |  | 34.44 | [5 empty] | +++ | GV+CA
346 | 22034070 | HOG _427 | Nguyen T.H | 1978 | 44 | 20.59 | [6 empty] | 20.01 |  | +++ | GV+MH
347 | 21050166 | HOG _428 | Bui N.O | 2001 | 21 | 16.82 |  |  | 27.63 | [5 empty] | +++ | GV+CA
348 | 22015893 | HOG _429 | Nguyen T.L | 2000 | 22 | 20.5 | [8 empty] | +++ | GV
349 | 18065574 | HOG _430 | Tran T.T | 1980 | 42 | 17.41 | [8 empty] | +++ | GV
350 | 13103010 | HOG _431 | Bui T.T.N | 1987 | 35 | 36.74 | [8 empty] | +++ | GV
351 | 22034031 | HOG _432 | Vu T.H | 1956 | 66 | 33.93 | [8 empty] | ++ | GV
352 | 22030572 | HOG _436 | Tran T.T.H | 1999 | 23 | 36.07 |  |  | 34.9 | [5 empty] | +++ | GV+CA
353 | 22034074 | HOG _437 | Nguyen T.H | 1990 | 32 | 16.97 |  |  | 31.89 | [5 empty] | +++ | GV+CA
354 | 22030638 | HOG _438 | Nguyen T.N.D | 1993 | 29 | 27.78 |  |  | 31.54 | [5 empty] | ++ | GV+CA
355 | 22033651 | HOG _439 | Le T.H.L | 2002 | 20 | 29.58 |  |  | 26.36 |  |  |  | 34.48 |  | +++ | GV+CA+MH
356 | 20073567 | HOG _443 | Pham T.T.T | 1991 | 31 | 30.45 | [8 empty] | +++ | GV
357 | 19057221 | HOG _444 | Nhu T.T.H | 1993 | 29 | 34.2 | [8 empty] | +++ | GV
358 | 22033638 | HOG _445 | Nguyen M.T | 2003 | 19 | 15.95 | [8 empty] | +++ | GV
359 | 18003035 | HOG _446 | Bui T.H | 2001 | 21 | 35.29 | [8 empty] | ++ | GV
360 | 15071544 | HOG _448 | Vu T.H | 1987 | 35 | 17.09 |  |  | 26.53 | [5 empty] | +++ | GV+CA
361 | 22033460 | HOG _449 | Nguyen T.T | 1971 | 51 | 25.87 |  |  |  | 26.16 |  | 36.62 | 22.85 |  | +++ | GV+TV+HSV-2+MH
362 | 22034124 | HOG _450 | Vu T.N.H | 1996 | 26 | 16.91 |  |  | 27.21 | [5 empty] | +++ | GV+CA
363 | 22030626 | HOG _451 | Nguyen K.H | 1975 | 47 | 14.45 | [8 empty] | +++ | GV
364 | 20024552 | HOG _452 | Nguyen T.T | 1993 | 29 | 37.53 | [8 empty] | +++ | GV
365 | 22033457 | HOG _453 | Loc K.V | 1996 | 26 | 21.32 |  |  | 27.44 | [5 empty] | ++ | GV+CA
366 | 22031054 | HOG _454 | Do T.D | 1993 | 29 | 15.62 | [8 empty] | ++ | GV
367 | 22059848 | HOG _455 | Hoang T.A | 1987 | 35 | 32.97 |  |  | 37.95 | [5 empty] | +++ | GV+CA
368 | 20014042 | HOG _456 | Nguyen T.H | 2000 | 22 | 15.14 | [8 empty] | +++ | GV
369 | 21006452 | HOG _457 | Nguyen T.T.H | 1991 | 31 | [9 empty] | ++ | Negative
370 | 22034189 | HOG _458 | Le T.K.T | 2001 | 21 | 18.12 | 35.7 |  | 27.41 | [5 empty] | ++ | GV+NG+CA
371 | 20082978 | HOG _459 | Dao M.H | 1998 | 24 | 29.8 |  | 36.6 | 28.36 | [5 empty] | +++ | GV+CT+CA
372 | 22035295 | HOG _460 | Bui H.G | 2002 | 20 | 16.83 |  |  | 29.92 | [5 empty] | ++ | GV+CA
373 | 13093496 | HOG _461 | Hoang T.K.O | 1976 | 46 | 17.47 |  |  | 29.76 | [5 empty] | +++ | GV+CA
374 | 22001314 | HOG _462 | Tran T.T.M | 2000 | 22 | 18.77 |  |  | 31.88 | [5 empty] | ++ | GV+CA
375 | 19095233 | HOG _463 | Doan P.T | 1999 | 23 | 28.59 |  | 36.88 | 28.8 | [5 empty] | +++ | GV+CT+CA
376 | 22034598 | HOG _464 | Tran T.P.N | 1998 | 24 | 19.28 | 35.99 |  | 27.96 | [5 empty] | +++ | GV+NG+CA
377 | 14036213 | HOG _465 | Nguyen T.P | 1990 | 32 | 14.54 |  | 36.37 | [6 empty] | +++ | GV+CT
378 | 12007269 | HOG _467 | Le N.A | 1988 | 34 | 15.6 |  |  | 28.53 | [5 empty] | ++ | GV+CA
379 | 22034632 | HOG _469 | Lo T.H | 1990 | 32 | 32.09 | 34.18 | 33.47 | [6 empty] | nomal | GV+NG+CT
380 | 22033348 | HOG _470 | Pham T.T.N | 1991 | 31 | 32.12 |  |  | 36.16 | [5 empty] | nomal | GV+CA
381 | 22021414 | HOG _471 | Ngo T.H | 1998 | 24 | 28.89 |  | 37.06 | [6 empty] | nomal | GV+CT
382 | 22034184 | HOG _472 | Nguyen T.P | 1983 | 39 | 27.19 | [8 empty] | nomal | GV
383 | 17105350 | HOG _473 | Le Q.M | 1990 | 32 | 28.68 | 36.21 | 23.91 | [6 empty] | nomal | GV+NG+CT
384 | 19053224 | HOG _474 | Pham T.T.H | 1998 | 24 | 18.53 |  |  | 28.52 | [5 empty] | +++ | GV+CA
385 | 22035276 | HOG _475 | Tran T.D.H | 1998 | 24 | 29.22 | [8 empty] | +++ | GV
386 | 22063330 | HOG _476 | Nguyen T.T | 1967 | 55 | 31.67 | 36.66 | 36.22 |  |  |  |  | 22.92 |  | +++ | GV+NG+CT+MH
387 | 17156838 | HOG _477 | Nguyen T.T.N | 1992 | 30 | 15.1 |  | 35.77 | 32.07 |  |  |  | 19.42 |  | +++ | GV+CT+CA+MH
388 | 21072558 | HOG _479 | Nguyen M.H | 1996 | 26 | 33.18 | [8 empty] | ++ | GV
389 | 22034703 | HOG _480 | Dao P.A | 1998 | 24 | 17.77 |  |  | 27.75 | [5 empty] | +++ | GV+CA
390 | 17022087 | HOG _482 | Do T.T | 1976 | 46 | 15.65 |  | 36.6 |  |  |  |  | 24.67 |  | +++ | GV+CT+MH
391 | 21034653 | HOG _483 | Bui T.T | 1973 | 49 | 15.03 |  |  | 27.07 | [5 empty] | +++ | GV+CA
392 | 15032466 | HOG _485 | Do T.T | 1988 | 34 | 31.97 |  |  |  | 24.18 |  |  |  |  | +++ | GV+TV
393 | 10112640 | HOG _486 | Nguyen T.M.L | 1985 | 37 | 14.69 |  |  | 27.9 | [5 empty] | +++ | GV+CA
394 | 12081133 | HOG _488 | Nguyen T.G | 1990 | 32 | 25.68 | [8 empty] | +++ | GV
395 | 22034722 | HOG _489 | Giang T.T | 1977 | 45 | 33.17 | [8 empty] | +++ | GV
396 | 20017656 | HOG _490 | Pham T.T | 1976 | 46 | 20.78 |  |  | 27.77 | [5 empty] | +++ | GV+CA
397 | 12099900 | HOG _491 | Nguyen T.H | 1976 | 46 | 15.48 | 35.13 | 34.25 | 30.74 | [5 empty] | ++ | GV+NG+CT+CA
398 | 12072939 | HOG _492 | Nguyen T.H | 1989 | 33 | 27.08 | [8 empty] | nomal | GV
399 | 22032216 | HOG _494 | Bui T.H | 1998 | 24 | 16.15 |  |  | 28.14 | [5 empty] | +++ | GV+CA
400 | 21040207 | HOG _495 | Nguyen T.H | 1991 | 31 | 15.96 |  |  | 30.52 | [5 empty] | +++ | GV+CA
401 | 22033682 | HOG _496 | Hoang H.N | 1996 | 26 | 36.18 |  |  | 32.85 | [5 empty] | +++ | GV+CA
402 | 22021667 | HOG _497 | Do P.Q | 1998 | 24 | 34.6 | [8 empty] | +++ | GV
403 | 22033675 | HOG _498 | Dinh T.D | 1998 | 24 | 29.94 |  |  | 29.62 | 23.27 |  |  |  |  | +++ | GV+CA+TV
404 | 19086045 | HOG _499 | Hoang A.P | 1993 | 29 | 32.31 |  |  | 31.97 | [5 empty] | ++ | GV+CA
405 | 21180893 | HOG _500 | Pham T.Y | 1993 | 29 | 14.86 |  |  | 31.88 | [5 empty] | +++ | GV
406 | 19011397 | HOG _501 | Nguyen T.L | 1996 | 26 | 15.14 | [8 empty] | +++ | GV
407 | 19068380 | HOG _502 | Pham T.T | 1992 | 30 | 32.47 | [8 empty] | +++ | GV
408 | 22047537 | HOG_503 | Nguyen T.T | 1988 | 34 | 30.23 | 20.92 | 32.03 | 27.07 |  |  |  | 16.56 |  | ++ | GV+NG+CT+CA+MH
409 | 11054030 | HOG _504 | Kim T.T | 1983 | 39 | 30.63 |  |  | 28.53 | [5 empty] | ++ | GV+CA
410 | 22035249 | HOG _505 | Duong T.S | 1997 | 25 | 27.41 |  |  | 26.77 | [5 empty] | ++ | GV+CA
411 | 22034774 | HOG _506 | Trinh T.T | 1987 | 35 | 15.63 |  |  | 27.89 | [5 empty] | ++ | GV+CA
412 | 22028146 | HOG _507 | Do T.T | 1995 | 27 | 32.64 | [8 empty] | +++ | GV
413 | 13047428 | HOG _508 | Tran T.H | 1985 | 37 | 30.99 | [6 empty] | 31.86 |  | +++ | GV+MH
414 | 22035025 | HOG _509 | Nguyen T.T | 1996 | 26 | 16.93 | [8 empty] | +++ | GV
415 | 17108962 | HOG _510 | Nguyen T.T,N | 1994 | 28 | 13.98 |  | 19.89 | 28.21 |  |  |  | 18.22 |  | +++ | GV+CT+CA+MH
416 | 22033721 | HOG _511 | Nguyen H.T | 1990 | 32 | 17.14 |  | 36.93 | 29.81 |  |  |  | 21.44 |  | ++ | GV+CT+CA+MH
417 | 22035231 | HOG _512 | Dinh T.A.T | 2001 | 21 | 35.06 |  |  | 25.62 | [5 empty] | +++ | GV+CA
418 | 22034880 | HOG _513 | Hoang T.P | 2000 | 22 |  |  |  | 29.36 |  |  |  | 31.35 |  | +++ | CA+MH
419 | 18039149 | HOG _514 | Phan D.M | 1995 | 27 |  |  |  | 31.77 | [5 empty] | ++ | CA
420 | 22035186 | HOG _515 | Bui T.M.P | 2000 | 22 | [7 empty] | 32.01 |  | +++ | MH
421 | 22035194 | HOG _516 | Vu T.H | 2001 | 21 |  |  |  | 27.03 | [5 empty] | +++ | CA
422 | 12149472 | HOG _517 | Do T.H | 1988 | 34 |  |  |  | 26.09 | [5 empty] | ++ | CA
423 | 8039769 | HOG _518 | Nguyen T.P | 1984 | 38 | 15.13 | [6 empty] | 24.56 |  | +++ | GV+MH
424 | 22034866 | HOG _519 | Hoang T.N.B | 1990 | 32 | 36.99 | [8 empty] | ++ | GV
425 | 21075053 | HOG _520 | Phan T.T.A | 1995 | 27 | 28.91 |  |  | 29.63 | [5 empty] | ++ | GV+CA
426 | 22035207 | HOG _521 | Nguyen H.N | 1991 | 31 | 28.93 | [8 empty] | ++ | GV
427 | 22033701 | HOG _522 | Le T.T | 1973 | 49 | 33.05 | [8 empty] | +++ | GV
428 | 19133756 | HOG _523 | Nong T.H | 1985 | 37 | 34.82 |  |  | 30.14 | [5 empty] | +++ | GV+CA
429 | 22035213 | HOG _524 | Nguyen T.H.T | 1993 | 29 | 35.71 |  |  | 30.54 | [5 empty] | +++ | GV+CA
430 | 22033726 | HOG _525 | Le T.TH | 1992 | 30 | 19.64 |  |  | 27.7 | [5 empty] | +++ | GV+CA
431 | 22033699 | HOG _526 | Le T.T.M | 1975 | 47 | 36.14 |  |  | 29.63 | [5 empty] | ++ | GV+CA
432 | 30201794 | HOG _527 | Pham T.P.T | 1979 | 43 | 21.93 |  |  | 32.12 | [5 empty] | +++ | GV+CA
433 | 22025852 | HOG _529 | Vũ T.N | 1993 | 29 | 23.43 | [8 empty] | nomal | GV
434 | 22033741 | HOG _530 | Nguyen T.H | 1996 | 26 | 33.64 | [8 empty] | nomal | GV
435 | 22038388 | HOG_531 | Dang T.V | 1998 | 24 | 22.97 |  |  | 34.02 | [5 empty] | ++ | GV+CA
436 | 21072350 | HOG_532 | Pham T.N.A | 1998 | 24 | 18.96 |  |  | 39.94 | [5 empty] | +++ | GV+CA
437 | 18008217 | HOG_533 | Nguyen T.T | 1984 | 38 | 16.5 |  |  | 31.04 | [5 empty] | ++ | GV+CA
438 | 12145236 | HOG_534 | Nguyen T.M | 1963 | 59 | 15.76 | [6 empty] | 36.07 |  | +++ | GV+MH
439 | 14126592 | HOG_535 | Pham T.D | 1992 | 30 | 13.98 |  |  | 26.85 | [5 empty] | + | GV+CA
440 | 18096348 | HOG_536 | Nguyen T.Q | 1991 | 31 |  |  |  | 39.01 | [5 empty] | ++ | CA
441 | 22038360 | HOG_537 | Tran T.M.H | 2000 | 22 | 25.75 |  |  | 33.67 | [5 empty] | +++ | GV+CA
442 | 17114065 | HOG_538 | Le T.B | 1984 | 38 | 34.95 | [8 empty] | ++ | GV+CA
443 | 22038347 | HOG_539 | Nguyen T.T | 1992 | 30 | 14.91 | [8 empty] | ++ | GV
444 | 15010530 | HOG_540 | Nguyen H.D | 1974 | 48 | 19.04 | [8 empty] | ++ | GV
445 | 16093501 | HOG_541 | Can T.V.T | 1995 | 27 | 40.68 |  | 24.48 | [6 empty] | ++ | GV+CT
446 | 22187352 | HOG_542 | Nguyen T.H | 1981 | 41 | 16.33 | [8 empty] | ++ | GV
447 | 22038346 | HOG_543 | Dang T.T | 1990 | 32 | 27.19 | [8 empty] | ++ | GV
448 | 13056763 | HOG_544 | Nguyen T.Q.N | 1988 | 34 | 33.61 | [8 empty] | ++ | GV
449 | 12119668 | HOG_545 | Le T.Q | 1990 | 32 | 26.37 |  | 37.03 | [6 empty] | ++ | GV+CT
450 | 22038019 | HOG_546 | Dao T.C | 1976 | 46 | 31.03 | [8 empty] | ++ | GV
451 | 7070304 | HOG_547 | Nguyen T.H | 1972 | 50 | 28.83 | [8 empty] | ++ | GV
452 | 10090816 | HOG_548 | Do K.N | 1988 | 34 | 26.03 |  | 35.02 | [6 empty] | ++ | GV+CT
453 | 16002897 | HOG_549 | Ngo T.T.M | 1994 | 28 | 14.88 |  |  | 25.74 | [5 empty] | ++ | GV+CA
454 | 20080946 | HOG_550 | Lua T.K.D | 2000 | 22 | 18.39 | [8 empty] | ++ | GV
455 | 18137260 | HOG_551 | Nguyen T.H | 1991 | 31 | 19.92 |  | 24.12 | 29.68 | [5 empty] | +++ | GV+CT+CA
456 | 21072123 | HOG_552 | Doan T.T | 1987 | 35 | 17.02 |  |  | 30.16 | [5 empty] | +++ | GV+CA
457 | 22042921 | HOG_553 | Phan N.M | 2004 | 18 | 25.41 |  | 36.43 | 34 | [5 empty] | ++ | GV+CT+CA
458 | 7104861 | HOG_554 | Hoa T.H.H | 1983 | 39 | 26.89 | [8 empty] | +++ | GV
459 | 22047983 | HOG_555 | Thai H.N | 1991 | 31 | 14.64 |  |  | 32.98 | [5 empty] | +++ | GV+CA
460 | 20047413 | HOG_556 | Dang T.L | 1995 | 27 | 16.19 |  |  | 30.21 |  |  |  | 24.65 |  | +++ | GV+CA+MH
461 | 21056713 | HOG_557 | Le T.H | 1990 | 32 | 26.59 |  |  | 33.62 | [5 empty] | ++ | GV+CA
462 | 20220392 | HOG_558 | Le P.L | 1992 | 30 | 15.22 |  |  | 31.03 | [5 empty] | +++ | GV+CA
463 | 16070819 | HOG_559 | Nguyen T.T | 1994 | 28 | 14.11 |  |  | 28.56 |  |  |  | 36.02 |  | +++ | GV+CA+MH
464 | 10091889 | HOG_560 | Nguyen T.T | 1965 | 57 | 24.61 | [8 empty] | +++ | GV
465 | 17078012 | HOG_561 | Nguyen T.K | 1987 | 35 | 35.44 | [8 empty] | +++ | GV
466 | 6900470 | HOG_562 | Ta T.T | 1973 | 49 | 15.31 |  |  | 30.3 | [5 empty] | +++ | GV+CA
467 | 22047964 | HOG_563 | Nguyen L.V | 2001 | 21 | 30.17 | [8 empty] | ++ | GV
468 | 22038278 | HOG_564 | Pham T.H | 1998 | 24 | 26.69 |  |  | 33.4 | [5 empty] | +++ | GV+CA
469 | 22047575 | HOG_565 | Nguyen T.K.O | 1999 | 23 | 17.18 | [8 empty] | +++ | GV
470 | 14108700 | HOG_566 | Nguyen T.Y | 1993 | 29 | 32.4 | [8 empty] | +++ | GV
471 | 17106317 | HOG_567 | Nguyen T.T.H | 1987 | 35 | 15.03 |  | 24.53 | 34.71 | [5 empty] | +++ | GV+CT+CA
472 | 22045536 | HOG_568 | Nguyen T.H | 1990 | 32 | 29.68 |  | 35.93 | [6 empty] | +++ | GV+CT
473 | 20097246 | HOG_569 | Nguyen T.T | 1992 | 30 | 17.25 |  | 36.5 | [6 empty] | +++ | GV+CT
474 | 17117041 | HOG_570 | Le T.L | 1971 | 51 | 26.45 |  | 36.21 | [6 empty] | ++ | GV+CT
475 | 22028432 | HOG_571 | Ninh T.P | 1985 | 37 | 16.58 |  |  | 26.23 | [5 empty] | ++ | GV+CA
476 | 22043132 | HOG_572 | Tran T.T | 1997 | 25 | 15.86 |  | 36.64 | 36.66 | [5 empty] | ++ | GV+CT+CA
477 | 22043131 | HOG_573 | Hoàng T.H.T | 2002 | 20 | 13.99 |  | 25.08 | [6 empty] | +++ | GV+CT
478 | 22046019 | HOG_574 | Nguyen T.T | 1997 | 25 | 37.12 | 37.58 |  | 28.72 | [5 empty] | +++ | GV+NG+CA
479 | 22038284 | HOG_575 | Ngo T.H | 1990 | 32 | 24.23 |  |  | 38.05 | 37.62 |  |  |  |  | ++ | GV+CA+TV
480 | 15150763 | HOG_576 | Ngo T.N | 1989 | 33 | 17.23 |  |  | 27.61 | [5 empty] | +++ | GV+CA
481 | 19652477 | HOG_577 | Chu H.P | 1984 | 38 | 14.89 |  | 37.09 | 38.66 | [5 empty] | +++ | GV+CT+CA
482 | 18159901 | HOG_578 | Tran T.S | 1989 | 33 |  | 37.12 |  | 26.24 | [5 empty] | ++ | NG+CA
483 | 19081894 | HOG_579 | Pham T.H | 1970 | 52 | 28.83 | [8 empty] | ++ | GV
484 | 17006206 | HOG_580 | Vu T.T.H | 1979 | 43 | 34.29 | [8 empty] | ++ | GV
485 | 22042954 | HOG_581 | Nguyen H.A | 1996 | 26 | 14.71 |  |  | 26.37 | [5 empty] | ++ | GV+CA
486 | 20088056 | HOG_582 | Truong T.H.Y | 1990 | 32 | 15.15 |  |  | 25.11 | [5 empty] | ++ | GV+CA
487 | 11054281 | HOG_583 | Le T.B.Y | 1982 | 40 | 16.52 |  |  |  | 24.29 |  |  | 23.09 | 24.63 | +++ | GV+TV+MH+MG
488 | 22045478 | HOG_584 | Hoang T.K.M | 2000 | 22 | 23.72 |  |  | 29.95 |  |  |  | 21.86 |  | ++ | GV+CA+MH
489 | 15086972 | HOG_585 | Nguyen T.V | 1994 | 28 | 35.66 |  |  | 26.58 | [5 empty] | ++ | GV+CA
490 | 18050180 | HOG_586 | Le T.A | 1997 | 25 | 36.13 | [8 empty] | ++ | GV
491 | 14138738 | HOG_587 | Quach T.T | 1989 | 33 | 34.44 | 36.46 |  | 25.66 | [5 empty] | +++ | GV+NG+CA
492 | 22022906 | HOG_588 | Nguyen L.D.H | 2003 | 19 | 33.85 |  |  | 28.58 | [5 empty] | +++ | GV+CA
493 | 22043009 | HOG_589 | Tran T.X | 1986 | 36 | 36.76 | [8 empty] | +++ | GV
494 | 22043011 | HOG_590 | Nguyen H.P | 1999 | 23 | 31.07 | [8 empty] | +++ | GV
495 | 22041020 | HOG_591 | Pham M.H | 1979 | 43 |  | 38.52 | [7 empty] | ++ | NG
496 | 22035252 | HOG_592 | Hoang T.L | 1988 | 34 | 31.67 | 22.53 | 33.93 |  |  |  | 18.58 |  | 28.32 | +++ | GV+NG+CT+HSV-2+MG
497 | 15136415 | HOG_593 | Tran T.L | 1987 | 35 |  |  |  | 30.02 | [5 empty] | ++ | CA
498 | 19123918 | HOG_594 | Do T.N | 1989 | 33 |  |  |  | 30.5 | [5 empty] | ++ | CA
499 | 16029868 | HOG_595 | Dao T.T.H | 1988 | 34 | 34.52 |  |  | 31.92 | [5 empty] | ++ | GV+CA
500 | 12135318 | HOG_596 | Nguyen T.T | 1992 | 30 | 35.78 | 37.2 | [7 empty] | ++ | GV+NG
501 | 22016444 | HOG_597 | Nguyen T.T.H | 1998 | 24 | 16.82 | 32.6 | 31.62 | 31.22 | [5 empty] | ++ | GV+NG+CT+CA
502 | 16025626 | HOG_598 | Tran T.T.H | 1982 | 40 | 16.77 | 29.53 | 29.86 | 31.26 | [5 empty] | ++ | GV+NG+CT+CA
503 | 17106727 | HOG_599 | Nguyen T.N | 1997 | 25 | 20.3 |  |  | 29.4 | [5 empty] | ++ | GV+CA
504 | 22047736 | HOG_600 | Ngyuyen T.T | 1995 | 27 | [9 empty] | +++ | Negative
505 | 22047722 | HOG_601 | Nguyen T.T.U | 1992 | 30 | 33.78 |  | 36.17 | 27.69 | [5 empty] | +++ | GV+CT+CA
506 | 22042248 | HOG_602 | Dao L.A | 2001 | 21 | 17.35 | 34.45 | 24.61 | 32.03 |  |  |  | 22.15 |  | ++ | GV+NG+CT+CA+MH
507 | 8070026 | HOG _503 | Dang N.A | 1984 | 38 | 16.84 | [8 empty] | +++ | GV
508 | 21024942 | HOG_604 | Nguyen T.Y | 1995 | 27 | 27.59 |  |  | 27.27 | [5 empty] | ++ | GV+CA
509 | 20046677 | HOG_605 | Tong T.N | 1999 | 23 | 16.79 |  |  | 31.65 |  | 18.45 |  | 37.61 |  | ++ | GV+CA+HSV-1+MH
510 | 22043015 | HOG_606 | Nguyen Q.L | 1999 | 23 | 18.85 | 36.07 | 36.05 | 31.64 | [5 empty] | ++ | GV+NG+CT+CA
511 | 6807796 | HOG_607 | Hoang T.V | 1984 | 38 | 22.71 |  | 32.38 | 29.6 | [5 empty] | ++ | GV+CT+CA
512 | 22042981 | HOG_608 | Tran T.H | 1990 | 32 | 15.07 |  | 35.47 | 32.46 | [5 empty] | ++ | GV+CT+CA
513 | 18279068 | HOG_609 | Tran T.D | 1997 | 25 |  | 29.64 |  | 31.34 | [5 empty] | ++ | NG+CA
514 | 22039861 | HOG_610 | Le T.H | 1988 | 34 | 27.61 | 29.81 | 30.33 | [6 empty] | ++ | GV+NG+CT
515 | 22043102 | HOG_611 | Be T.L | 1986 | 36 | 13.83 |  |  | 24.95 | [5 empty] | +++ | GV+CA
516 | 22047750 | HOG_612 | Truong T.T.H | 1982 | 40 | 34.9 |  |  | 30.99 | [5 empty] | ++ | GV+CA
517 | 19100247 | HOG_613 | Dinh T.H | 1988 | 34 | 31.17 |  |  | 28.67 |  |  |  | 32.68 |  | +++ | GV+CA
518 | 19009586 | HOG_614 | Vu T.H.T | 1996 | 26 | 19.72 |  |  |  | 21.67 |  |  | 16.34 |  | +++ | GV+TV+MH
519 | 22038019 | HOG_615 | Dao T.C | 1976 | 46 | 36.01 | [8 empty] | ++ | GV
520 | 20042645 | HOG_616 | Nguyen T.T.H | 1999 | 23 | 34.62 | [8 empty] | ++ | GV
521 | 22038314 | HOG_617 | Tran T.T.V | 1972 | 50 | 34.7 |  |  | 39.58 | [5 empty] | ++ | GV+CA
522 | 22038319 | HOG_618 | Tran T.T.H | 1969 | 53 | 15.83 |  |  | 26.37 | [5 empty] | +++ | GV+CA
523 | 22025043 | HOG_619 | Truong Q.A | 2001 | 21 | 14.96 |  |  | 26.43 | [5 empty] | ++ | GV+CA
524 | 19090688 | HOG_620 | Phung T.T.H | 1994 | 28 | 14.13 |  |  | 31.39 | [5 empty] | +++ | GV+CA
525 | 22045385 | HOG_621 | Nguyen T.N | 1990 | 32 | 18.76 | [8 empty] | +++ | GV
526 | 8052430 | HOG_622 | Trinh T.T | 1988 | 34 | 15.59 | [8 empty] | +++ | GV
527 | 12003959 | HOG_623 | Do T.H | 1988 | 34 | 25.84 |  |  | 27.87 | [5 empty] | ++ | GV+CA
528 | 21021555 | HOG_624 | Nguyen T.H | 1992 | 30 | 15.34 | 38.3 |  | 28.91 | [5 empty] | +++ | GV+NG+CA
529 | 19096752 | HOG_625 | Doan H.G | 1994 | 28 | 34.29 |  | 39.72 | [6 empty] | +++ | GV+CT
530 | 22045484 | HOG_626 | Hoang T.H.Y | 1982 | 40 | 36.85 | 38.08 | 38.62 | [6 empty] | +++ | GV+NG+CT
531 | 19151373 | HOG_627 | Do T.H | 2000 | 22 | 19.17 |  |  | 29.54 |  |  |  | 23.66 |  | +++ | GV+CA+MH
532 | 31002914 | HOG_628 | Luong T.L | 1977 | 45 | 15.4 |  |  | 28.45 | [5 empty] | +++ | GV+CA
533 | 17133050 | HOG_629 | Nguyen T.M.A | 1995 | 27 | 15.14 |  |  | 32.28 |  |  |  | 30.76 |  | +++ | GV+CA+MH
534 | 21070292 | HOG_630 | Vu T.T.H | 1996 | 26 | 17.58 | [8 empty] | ++ | GV
535 | 22045386 | HOG_631 | Nguyen T.L.A | 1977 | 45 | 34.08 | 36.65 | 38.14 | [6 empty] | +++ | GV+NG+CT
"Non": Negative
Demographic characteristics
|  |  | Age | n | %
|  |  | Min | 18 | 3.364%
|  |  | Max | 66 | 12.34%
|  |  | 18-49 | 510 | 95.33%
|  |  | >49 | 25 | 4.673%
|  |  | Total | 535 | 100%
|  |  | Leukocyte densities | n | %
|  |  | Nomal | 31 | 5.794%
|  |  | "+" (<5 leu/microsphere) | 2 | 0.3738%
|  |  | "++" (5-10 leu/microsphere) | 181 | 33.83%
|  |  | "+++" (>10 leu/microsphere) | 321 | 60%
|  |  | "++" & "+++" (>5 leu/microsphere) | 502 | 93.83%
|  |  | Total | 535 | 100%
[6 empty] | 13.55 |  |  | 26.25
[6 empty] | 44909 |  |  | 44616